# Software Version 6.7

# DICOM Conformance Statement

## Related Products:

PowerServer, Gateway, Gateway Router, PowerCache, RapidResults

All Information provided in this document and in the accompanying software is subject to change without notice and does not represent a commitment on the part of RamSoft. RamSoft assumes no responsibility for any errors or consequential damages that may result from the use or misinterpretation of any information in this document. No part of this document may be reproduced in any form or by any means, electronic or mechanical, including photocopying and recording, for any purpose without the express written permission of RamSoft.

MANUFACTURER

RamSoft Incorporated

90 Adelaide St W, Suite #900 Toronto, ON M5H 3V9 CANADA

Toll Free:	(888) 343-9146

Fax:	(416) 674-7147

Web Site:	http://www.ramsoft.com

RamSoft®, RamSoft PACS™ are registered trademark of RamSoft Inc.

PowerServer™, PowerReader™, PowerCache™, Essence™, RapidResults™ and Gateway™ are registered

trademarks of RamSoft Inc.

Microsoft® and Windows® are registered trademarks of Microsoft Corporation. M*Modal Speech Understanding™ is a registered trademark of MModal IP LLC. Dragon Medical® is a registered trademark of Nuance Communications, Inc.

# CONFORMANCE STATEMENT OVERVIEW

The DICOM 3.0 Conformance Statement for the software produced by RamSoft including PowerServer, Gateway, Gateway Router, PowerCache and RapidResults

PowerServer and Gateway are self-contained network computer systems used for capturing, storing, displaying, reporting, transmitting and archiving diagnostic medical images. The system conforms to the DICOM

3.0 standard to share medical data with other medical imaging devices.

#### Table 1-1 Network Services

| SOP Classes | User of Service (SCU) | Provider of Service (SCP) |
| --- | --- | --- |
| Transfer |  |  |
| Computed Radiography Image Storage | Stored and Viewed | Yes |
| CT Image Storage | Stored and Viewed | Yes |
| Enhanced CT Image Storage | Stored and Viewed | Yes |
| PET Image Storage | Stored and Viewed | Yes |
| PET Curve Storage | Stored and Viewed | Yes |
| Enhanced PET Image Storage | Stored and Viewed | Yes |
| Digital X-Ray Image Storage for Presentation | Stored and Viewed | Yes |
| Digital X-Ray Image Storage for Processing | Stored Viewing is Configurable | Yes |
| Digital Mammography X-Ray Image Storage for Presentation | Stored and Viewed | Yes |
| Digital Mammography X-Ray Image Storage for Processing | Stored Viewing is Configurable | Yes |
| Breast Tomosynthesis Image Storage | Stored and Viewed | Yes |
| X-Ray Angiographic Image Storage | Stored and Viewed | Yes |

| SOP Classes | User of Service (SCU) | Provider of Service (SCP) |
| --- | --- | --- |
| Enhanced XA Image Storage | Stored and Viewed | Yes |
| RETIRED X-Ray Angiographic Bi-Plane Image Storage | Stored and Viewed | Yes |
| X-Ray Fluoroscopy Image Storage | Stored and Viewed | Yes |
| Enhanced XRF Image Storage | Stored and Viewed | Yes |
| Digital Intra Oral X-ray Image Storage for Presentation | Stored and Viewed | Yes |
| Digital Intra Oral X-ray Image Storage for Processing | Stored Viewing is Configurable | Yes |
| RT Image Storage | Stored and Viewed | Yes |
| MR Image Storage | Stored and Viewed | Yes |
| Enhanced MR Image Storage | Stored and Viewed | Yes |
| MR Spectroscopy Image Storage | Stored only | Yes |
| Nuclear Medicine Image Storage | Stored and Viewed | Yes |
| RETIRED Nuclear Medicine Image Storage | Stored and Viewed | Yes |
| Ultrasound Image Storage | Stored and Viewed | Yes |
| Ultrasound Multi-frame Image Storage | Stored and Viewed | Yes |
| RETIRED Ultrasound Image Storage | Stored and Viewed | Yes |
| RETIRED Ultrasound Multi Frame Image Storage | Stored and Viewed | Yes |
| Secondary Capture Image Storage | Stored and Viewed | Yes |
| Multi Frame Single Bit Secondary Capture Image Storage | Stored and Viewed | Yes |
| Multi Frame Grayscale Byte Secondary Capture Image Storage | Stored and Viewed | Yes |
| Multi Frame Grayscale Word Secondary Capture Image Storage | Stored and Viewed | Yes |
| Multi Frame True Color Secondary Capture Image Storage | Stored and Viewed | Yes |
| Hardcopy Grayscale Image Storage | Stored and Viewed | Yes |

| SOP Classes | User of Service (SCU) | Provider of Service (SCP) |
| --- | --- | --- |
| Hardcopy Color Image Storage | Stored and Viewed | Yes |
| Ophthalmic Photography 16 Bit Image Storage | Stored and Viewed | Yes |
| Ophthalmic Photography 8 Bit Image Storage | Stored and Viewed | Yes |
| RETIRED VL Image Storage | Stored and Viewed | Yes |
| VL Endoscopic Image Storage | Stored and Viewed | Yes |
| VL Microscopic Image Storage | Stored and Viewed | Yes |
| VL Photographic Image Storage | Stored and Viewed | Yes |
| VL Slide Coordinates Microscopic Image Storage | Stored and Viewed | Yes |
| RETIRED VL Multi Frame Image Storage | Stored and Viewed | Yes |
| Grayscale Softcopy Presentation State Storage | Stored and Viewed | Yes |
| Encapsulated PDF Storage | Stored and Viewed | Yes |
| Basic Text SR | Stored and Viewed | Yes |
| Enhanced SR | Stored and Viewed | Yes |
| Comprehensive SR | Stored and Viewed | Yes |
| Mammography CAD SR | Stored and Viewed | Yes |
| X-Ray Radiation Dose SR Storage | Stored only | Yes |
| Raw Data Storage | Stored only | Yes |
| Key Object Selection | Yes – Images only | Yes |
| Basic Voice Audio Waveform Storage | Stored and Listened | Yes |
| Storage Commitment Push Model | Yes | Yes |
| Fuji Computed Radiography Image Storage | Stored and Viewed | Yes |
| Query/Retrieve |  |  |

| SOP Classes | User of Service (SCU) | Provider of Service (SCP) |
| --- | --- | --- |
| Study Root Information Model - FIND | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Study Root Information Model - MOVE | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Study Root Information Model - GET | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Patient Root Information Model - FIND | No | Yes – Only hierarchical queries. |
| Patient Root Information Model - MOVE | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Patient Root Information Model - GET | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Patient/Study Information Model - FIND | No | Yes – Only hierarchical queries. |
| Patient/Study Only Information Model - MOVE | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Patient/Study Only Information Model - GET | Yes – Only hierarchical queries. | Yes – Only hierarchical queries. |
| Workflow Management |  |  |
| Modality Performed Procedure Step SOP Class | Yes | Yes |
| Modality Worklist Information Model - FIND | Yes | Yes |
| Print Management |  |  |
| Basic Grayscale Print Management Meta SOP Class | Yes | No |
| Basic Color Print Management Meta SOP Class | Yes | No |
| WADO |  |  |
| WADO – URL – Retrieve Rendered Imaging Document | No | Yes |

Table 1-2 Media Services

Note: Type of media supported on a given workstation depends on the peripherals connected to that workstation. Conformance is based on software support for media only.

| Media Storage Application Profile | Application Profile | Identifier | Write Files (FSC or FSU) | Read Files (FSR) |
| --- | --- | --- | --- | --- |
| Compact Disk – Recordable |  |  |  |  |
| General Purpose Interchange | General Purpose CD-R Interchange | STD-GEN-CD | Yes | Yes |
| Basic Cardiac | Basic Cardiac X-Ray Angiographic Studies on CD-R Media | STD-XABC-CD | No | Yes |
| 1024 X-Ray Angiographic | 1024 X-Ray Angiographic Studies on CD-R Media | STD-XA1K-CD | No | Yes |
| Ultrasound | Image Display for CD-R | STD-US-ID-MF-CDR | No | Yes |
| CT/MR | CT/MR Studies on CD-R | STD-CTMR-CD | No | Yes |
| DVD |  |  |  |  |
| General Purpose Interchange | General Purpose Interchange on DVD-RAM Media | STD-GEN-DVD-RAM | Yes | Yes |
| General Purpose Interchange with JPEG | General Purpose DVD Interchange with JPEG | STD-GEN-DVD-JPEG | No | Yes |
| General Purpose Interchange with JPEG 2000 | General Purpose DVD Interchange with JPEG 2000 | STD-GEN-DVD-J2K | No | Yes |
| 1024 X-Ray Angiographic | 1024 X-Ray Angiographic Studies on DVD Media | STD-XA1K-DVD | No | Yes |
| Ultrasound | Image Display for DVD | STD-US-ID-MF-DVD | No | Yes |
| CT/MR | CT/MR Studies on DVD Media | STD-CTMR-DVD | No | Yes |
| DVD-RAM |  |  |  |  |
| General Purpose Interchange | General Purpose Interchange on DVD-RAM Media | STD-GEN-DVD-RAM | No | Yes |

| Ultrasound | Image Display | STD-US-ID-MF-DVD | No | Yes |
| --- | --- | --- | --- | --- |
| CT/MR | CT/MR Studies on DVD-RAM Media | STD-CTMR-DVD-RAM | No | Yes |
| Blu-ray Disc |  |  |  |  |
| General Purpose Interchange | General Purpose Interchange on BD Media | STD-GEN-BD | Yes | Yes |
| General Purpose Interchange with JPEG | General Purpose BD Interchange with JPEG | STD-GEN-BD-JPEG | No | Yes |
| General Purpose Interchange with JPEG 2000 | General Purpose BD Interchange with JPEG 2000 | STD-GEN-BD-J2K | No | Yes |
| USB |  |  |  |  |
| General Purpose Interchange with JPEG | General Purpose USB Media Interchange with JPEG | STD-GEN-USB-JPEG | No | Yes |
| General Purpose Interchange with JPEG 2000 | General Purpose USB Media Interchange with JPEG-2000 | STD-GEN-USB-J2K | No | Yes |
| MultiMedia Card |  |  |  |  |
| General Purpose Interchange with JPEG | General Purpose MultiMedia Card Interchange with JPEG | STD-GEN-MMC-JPEG | No | Yes |
| General Purpose Interchange with JPEG 2000 | General Purpose MultiMedia Card Interchange with JPEG-2000 | STD-GEN-MMC-J2K | No | Yes |
| CompactFlash |  |  |  |  |
| General Purpose Interchange with JPEG | General Purpose CompactFlash Interchange with JPEG | STD-GEN-CF-JPEG | No | Yes |
| General Purpose Interchange with JPEG 2000 | General Purpose CompactFlash Interchange with JPEG-2000 | STD-GEN-CF-J2K | No | Yes |
| Digital Card |  |  |  |  |
| General Purpose Interchange with JPEG | General Purpose Digital Card Interchange with JPEG | STD-GEN-SD-J2K | No | Yes |

| General Purpose Interchange with JPEG 2000 | General Purpose Digital Card Interchange with JPEG-2000 | STD-GEN-SD-J2K | No | Yes |
| --- | --- | --- | --- | --- |

# TABLE OF CONTENTS

Contents

Related Products:	1

1	CONFORMANCE STATEMENT OVERVIEW	2

2	TABLE OF CONTENTS	9

3	INTRODUCTION	1

3.1	Revision History	1

3.2	Audience	1

3.3	Remarks	1

3.4	Terms and Definitions	1

3.5	Basics of DICOM Communication	1

3.6	Abbreviations	2

3.7	References	2

4	NETWORKING	3

4.1	Implementation Model	3

4.1.1	Application Data Flow	3

4.1.2	Functional Definitions of AE Titles	5

4.1.2.1	ECHO-SCP	5

4.1.2.2	ECHO-SCU	5

4.1.2.3	STORAGE-SCP	6

4.1.2.4	STORAGE-SCU	6

4.1.2.5	STORAGE-SCU-COMMITMENT	6

4.1.2.6	STORAGE-SCP-COMMITMENT	6

4.1.2.7	FIND-SCU	6

4.1.2.8	FIND-SCP	6

4.1.2.9	MOVE-SCU	6

4.1.2.10	GET-SCU	6

4.1.2.11	MOVE-SCP	7

4.1.2.12	GET-SCP	7

4.1.2.13	MWL-SCU	7

4.1.2.14	MWL-SCP	7

4.1.2.15	PRINT-SCU	7

4.1.3	Sequencing of Real-World Activities	7

4.2	RamSoft PACS AE Specification	8

4.2.1	ECHO-SCP	8

4.2.1.1	SOP CLASSES	8

4.2.2	ECHO-SCU	10

4.2.2.2	SOP CLASSES SUPPORTED BY ECHO-SCU	10

4.2.2.4	NUMBER OF ASSOCIATIONS FOR ECHO-SCU	10

4.2.3	STORAGE-SCP	12

4.2.3.2	SOP CLASSES SUPPORTED BY STORAGE-SCP AND STORAGE-SCU	12

4.2.3.4	MAXIMUM PDU SIZE FOR STORAGE-SCP	14

4.2.3.7	PRESENTATION CONTEXTS FOR STORAGE-SCP AND STORAGE-SCU	16

4.2.4	STORAGE-SCU	19

4.2.4.1	SOP CLASSES	19

4.2.5	STORAGE-SCU-COMMITMENT	25

4.2.6	STORAGE-SCP-COMMITMENT	26

4.2.7	FIND-SCP	27

4.2.7.2	Transfer Syntax Support	27

4.2.7.6	ACCEPTABLE AND PROPOSED PRESENTATION CONTEXTS FOR FIND-SCP AND FIND-SCU	29

4.2.7.7	STUDY ROOT REQUEST IDENTIFIER FOR FIND-SCP	35

4.2.7.8	RESPONSE STATUS FOR FIND-SCP AND RECEIVE QUERY REQUEST	38

4.2.8	FIND-SCU	39

4.2.8.2	Transfer Syntax Support	39

4.2.8.6	RESPONSE STATUS FOR FIND-SCU AND QUERY REMOTE AE REQUEST	42

4.2.9	MOVE-SCP	43

4.2.9.2	Transfer Syntax Support	44

4.2.9.6	PATIENT ROOT AND PATIENT / STUDY ONLY REQUESTS IDENTIFIER FOR MOVE-SCP	48

4.2.9.7	RESPONSE STATUS FOR MOVE-SCP AND SEND TO REMOTE QUERY REQUEST	49

4.2.10	MOVE-SCU	51

4.2.10.2	Transfer Syntax Support	51

4.2.10.4	NUMBER OF ASSOCIATIONS AS A SCP FOR FIND-SCU	52

4.2.11	GET-SCP and GET-SCU	54

4.2.11.2	SOP CLASSES SUPPORTED BY GET-SCP AND GET-SCU	54

4.2.11.3	Transfer Syntax Support	55

4.2.12	MWL-SCP	57

4.2.12.3	NUMBER OF ASSOCIATIONS AS A SCP FOR MWL-SCP	57

4.2.12.6	SET OF RETURNED ATTRIBUTES FOR MWL-SCP	59

4.2.12.7	SET OF RETURNED ATTRIBUTES FOR MWL-SCP	59

4.2.13	MWL-SCU	62

4.2.14	STORAGE-SCP-MPPS and STORAGE-SCU-MPPS	64

4.2.14.2	SOP CLASSES SUPPORTED BY STORAGE-SCP-MPPS AND STORAGE-SCU-MPPS	64

4.2.15	PRINT-SCU	65

4.2.15.2	SOP CLASSES SUPPORTED BY PRINT-SCU	65

4.2.15.4	MAXIMUM PDU SIZE SENDED FOR PRINT-SCU	65

4.2.15.5	NUMBER OF ASSOCIATIONS FOR PRINT-SCU	65

4.2.15.7	PROPOSED PRESENTATION CONTEXTS FOR PRINT-SCU	66

4.2.16	WADO Specifications	68

4.2.17	DICOM WEB Support	71

4.3	Network Interfaces	77

4.3.1	Physical Network Interface	77

4.3.2	IPv4 and IPv6 support	77

4.4	Configuration	77

4.4.1	AE Title/Presentation Address Mapping	77

4.4.2	The main interfaces for configuration settings	77

4.4.3	Parameters	87

5	MEDIA INTERCHANGE	90

5.1	Implementation Model	90

5.1.1	Application Data Flow	90

5.1.2	Functional Definition of AEs	90

5.1.3	Sequencing of Real-Word Activities	90

5.1.4	File Meta Information Options	91

5.2	AE Specifications	91

5.2.1	Media Application Entity Specification	91

6	TRANSFORMATION OF DICOM TO CDA	92

7	SUPPORT OF CHARACTER SETS	92

7.1	Character Sets	92

8	SECURITY	94

8.1	Security Profiles	94

8.2	Association Level Security	96

8.3	Application Level Security	96

9	ANNEXES	97

9.1	IOD Contents	97

9.1.1	Created SOP Instances	97

9.1.2	Usage of Attributes From Received IODs	135

9.1.3	Attribute Mapping	148

9.1.4	Coerced / Modified Fields	155

9.2	Data Dictionary of Private Attributes	156

10	DICOM Patient, Visit, Study and Series Fields stored in DB	165

11	Revision History	167

# INTRODUCTION

## Revision History

Revision History is listed in Chapter 10 for standardization with other RamSoft Documents.

## Audience

This document is written for those who need to understand how RamSoft products will integrate into their healthcare facility. This includes both those responsible for overall imaging network policy and architecture, as well as integrators who need to have a detailed understanding of the DICOM features of the product. This document contains some basic DICOM definitions so that any reader may understand how this product implements DICOM features. However, integrators are expected to fully understand all the DICOM terminology, how the tables in this document relate to the product’s functionality, and how that functionality integrates with other devices that support compatible DICOM features

## Remarks

The scope of this DICOM Conformance Statement is to facilitate integration between RamSoft products and other DICOM products. The Conformance Statement should be read and understood in conjunction with the DICOM Standard. DICOM by itself does not guarantee interoperability. The Conformance Statement does, however, facilitate a first-level comparison for interoperability between different applications supporting compatible DICOM functionality.

This Conformance Statement is not supposed to replace validation with other DICOM equipment to ensure proper exchange of intended information. In fact, the user should be aware of the following important issues:

- The comparison of different Conformance Statements is just the first step towards assessing interconnectivity and interoperability between the product and other DICOM conformant equipment.

- Test procedures should be defined and executed to validate the required level of interoperability with specific compatible DICOM equipment, as established by the healthcare facility.

RamSoft PowerServer has participated in an industry-wide testing program sponsored by Integrating the Healthcare Enterprise (IHE). The IHE Integration Statement for RamSoft PowerServer, together with the IHE Technical Framework, may facilitate the process of validation testing.

## Terms and Definitions

This conformance Statement uses terms and definition from Chapter A.3.4 of Standard Digital Imaging and Communications in Medicine Part 2: Conformance.

## Basics of DICOM Communication

This section describes terminology used in this Conformance Statement for the non-specialist. The key terms used in the Conformance Statement are highlighted in italics below. This section is not a substitute for training about DICOM, and it makes many simplifications about the meanings of DICOM terms.

Two Application Entities (devices) that want to communicate with each other over a network using DICOM protocol must first agree on several things during an initial network “handshake”. One of the two devices must initiate an Association (a connection to the other device), and ask if specific services, information, and encoding can be supported by the other device (Negotiation).

DICOM specifies a number of network services and types of information objects, each of which is called an Abstract Syntax for the Negotiation. DICOM also specifies a variety of methods for encoding data, denoted Transfer Syntaxes. The Negotiation allows the initiating Application Entity to propose combinations of Abstract Syntax and Transfer Syntax to be used on the Association; these combinations are called Presentation Contexts. The receiving Application Entity accepts the Presentation Contexts it supports.

For each Presentation Context, the Association Negotiation also allows the devices to agree on Roles – which one is the Service Class User (SCU - client) and which is the Service Class Provider (SCP - server). Normally the device initiating the connection is the SCU, i.e., the client system calls the server, but not always.

The Association Negotiation finally enables exchange of maximum network packet (PDU) size, security information, and network service options (called Extended Negotiation information).

The Application Entities, having negotiated the Association parameters, may now commence exchanging data. Common data exchanges include queries for worklists and lists of stored images, transfer of image objects and analyses (structured reports), and sending images to film printers. Each exchangeable unit of data is formatted by the sender in accordance with the appropriate Information Object Definition, and sent using the negotiated Transfer Syntax. There is a Default Transfer Syntax that all systems must accept, but it may not be the most efficient for some use cases. Each transfer is explicitly acknowledged by the receiver with a Response Status indicating success, failure, or that query or retrieve operations are still in process.

Two Application Entities may also communicate with each other by exchanging media (such as a CD-R). Since there is no Association Negotiation possible, they both use a Media Application Profile that specifies “pre-negotiated” exchange media format, Abstract Syntax, and Transfer Syntax.

## Abbreviations

This conformance Statement uses abbreviations from Chapter A.3.6 of Standard Digital Imaging and Communications in Medicine Part 2: Conformance.

## References

- NEMA PS2 Digital Imaging and Communications in Medicine (DICOM) Standard, available free at http://medical.nema.org/

- NEMA PS3 Digital Imaging and Communications in Medicine (DICOM) Standard, available free at http://medical.nema.org/

# NETWORKING

## Implementation Model

## Application Data Flow

RamSoft workstations and servers can send and receive DICOM data. DICOM parameters are configurable through the Station List. DICOM communications can be secured through TLS.

The RamSoft DICOM Service is typically installed to start up automatically. Once the computer is turned on and Windows is started, the RamSoft DICOM Service is ready for DICOM communication. The RamSoft PACS application is typically installed to start up once an authorized user logs into Windows. RamSoft client workstations connected directly to a RamSoft server do not require the RamSoft DICOM Service running and do not connect through DICOM to the RamSoft server.

A RamSoft workstation can query a DICOM Modality Work list server to obtain patient and study information. A third party workstation can query a RamSoft server to obtain patient, study, series and image information.

A RamSoft workstation may contain an option to capture images through a video frame grabber, film digitizer, or document scanner. Alternately, images may be imported from various digital formats. Images are stored on the hard drive.

DICOM data can be pushed from the server to a DICOM AE by clicking on the Push button and selecting Station for Target. Once Target and Destination stations are selected and the priority level is set, then data is sent in the background. If any failures occur, the error is recorded in the Transmit List. RamSoft PACS will retry transmission until it completes.

DICOM data can also be set to be automatically routed from a RamSoft server.

DICOM data can be requested from a DICOM AE through Query/Retrieve Service, by selecting Open or Request after finding the desired patients or studies on a remote workstation/server. Studies previously sent to an archive, may be retrieved automatically when accessed or retrieved from the Patient Explorer by clicking on the “Open data from server” icon.

DICOM data may be recorded on floppy disks, CD-R, CD-RW, or DVD media.

RamSoft DICOM Server can create, send and receive images, presentation states and SR document objects. It contains bi-directional query capabilities. RamSoft users can query and request information from other PACS stations. Many RamSoft workstations and all RamSoft servers can process query requests. Scheduling information can be queried from a DICOM Modality Worklist. RamSoft DICOM Server can format images for printing on a wide variety of DICOM printing devices. RamSoft DICOM Server can create, read and update archival filesets on floppy disks, CD-Rs, CD-RWs, DVDs, and other media.

#### Figure 4.1.1 Implementation Model

User requests

Transmission

Image stored on

hard drive

User queries HIS

worklist

User queries

PACS and requests images

Database is

queried and images are sent

User requests

printing of images

RamSoft AE

Images stored

on external system

External system

requests transmission

External system

searches its worklist

External system

searches its DB and	sends images

External system

queries	PACS and requests

External system

prints images

DICOM 3.0 Interface

RamSoft DICOM Server contains a single Application Entity that implements the Verification Service Class, Storage Service Class, Grayscale Softcopy Presentation State Storage SOP Class, Structured Reporting Storage SOP Class, Basic Worklist Management Service, the Query/Retrieve Service Class as (SCU, SCP), Modality Performed Procedure Step SOP Classes (SCU, SCP) and Storage Commitment Service Class (SCU,SCP). RamSoft PACS implements the Print Management Service Class as an SCU. RamSoft PACS implements the Media Storage Service Class as a File Set Creator (FSC), File Set Reader (FSR) and File Set Updater (FSU). Verification (SCU and SCP), Storage (SCU and SCP), Grayscale Softcopy Presentation State Storage (SCP), Structured Reporting Storage (SCP), Basic Worklist Management (SCU and SCP), Query (SCP)/Retrieve (SCU and SCP), Modality Performed Procedure Step (SCU and SCP) and Storage Commitment (SCU and SCP) are handled by the RamSoft DICOM Service. Other services are handled directly by the PowerServer application. All DICOM operations are performed by a single RamSoft PACS AE. The PACS AE can be modeled as the following AE Titles:

- ECHO-SCP, which responds to verification requests

- ECHO-SCU, which sends a verification request

- STORAGE-SCP, which receives incoming composite instances

- STORAGE-SCU, which sends outbound composite instances

- STORAGE-SCU-COMMITMENT, which manages verification of sending of composite instances

- STORAGE-SCP-COMMITMENT, which manages verification of receiving of composite instances

- FIND-SCP, which receives incoming queries for lists of studies

- FIND-SCU, which queries remote AEs for lists of studies

- GET-SCP and MOVE-SCP, which respond to requests for studies

- GET-SCU and MOVE-SCU, which retrieve selected studies

- MWL SCP, which receives incoming queries for worklists of patients

- MWL SCU, which queries remote AEs for lists worklists of patients

- STORAGE-SCU-MPPS, which manages sending of composite instances

- STORAGE-SCP-MPPS, which manages receiving of composite instances

- PRINT SCU, which prints requested instances to remote DICOM Printers

RamSoft DICOM AE support the Media Storage Application Profiles shown in the Table 1-2.

## Functional Definitions of AE Titles

### ECHO-SCP

ECHO-SCP waits in the background for connections, will accept associations with Presentation Contexts for SOP Class of the Verification Service Class, and will respond successfully to echo requests.

### ECHO-SCU

ECHO-SCU is activated through the user interface when a user selects a remote AE to verify (from the Station List), then initiates a “Test Connection.”

### STORAGE-SCP

STORAGE-SCP waits in the background for connections, will accept associations with Presentation Contexts for SOP Classes of the Storage Service Class, and will store the received instances to the local database where they may subsequently be listed and viewed through the user interface.

### STORAGE-SCU

STORAGE-SCU is activated through the user interface when a user selects instances from the local database, or the currently displayed instance, and requests that they be sent to a remote AE (selected from the Station List).

### STORAGE-SCU-COMMITMENT

STORAGE-SCU-COMMITMENT is activated through the user interface as an option in the Station List. If enabled for the remote AE, at the end of DICOM Association for STORAGE-SCU, N-ACTION is issued, to confirm the end of image transmission.

### STORAGE-SCP-COMMITMENT

STORAGE-SCP-COMMITMENT may be optionally used in conjunction with STORAGE-SCP. In addition to STORAGE-SCP, it receives N-ACTION to verify if all images were successfully stored. It sends N-EVENT-REPORT to STORAGE-SCU-COMMITMENT.

### FIND-SCU

FIND-SCU is activated through the user interface when a user selects a remote AE to query (from the Station List), then initiates a query. Queries are performed from the study through the series and instance levels until all matching instances have been listed.

### FIND-SCP

FIND-SCP waits in the background for connections, will accept associations with Presentation Contexts for SOP Classes of the FIND Service Classes, and will respond to query requests from FIND-SCU.

### MOVE-SCU

For each remote AE, either MOVE-SCU or GET-SCU may be configured in the Station List. MOVE-SCU may be activated through the user interface when a user selects a study, series or instance for retrieval. A connection to the remote AE is established to initiate and monitor the retrieval and the STORAGE-SCP AE receives the retrieved instances.

### GET-SCU

For each remote AE, either MOVE-SCU or GET-SCU may be configured in the Station List. The functional behavior of GET-SCU is identical to MOVE-SCU. The difference is that GET-SCU uses C-GET and images are received in the same DICOM association as the C-GET request. GET-SCU will work even if the SCU computer uses a dynamic IP address or hostname whereas MOVE-SCU will not. GET-SCU is also faster than MOVE-SCU, so it is recommended to use GET-SCU if supported by the remote AE.

### MOVE-SCP

MOVE-SCP receives requests from remote AEs to retrieve images. MOVE-SCP accepts an association from STORE-SCU to transfer the requested images.

### GET-SCP

The functional behavior of GET-SCP is identical to MOVE-SCP. The difference that GET-SCP uses C-GET request And C-GET response to transfer the requested images using the same DICOM association as the C-GET request

### MWL-SCU

MWL-SCU is activated through the user interface when a user selects Advanced Search ->Remote Worklist and remote station.

### MWL-SCP

FIND-SCP waits in the background for connections, will accept associations with Presentation Contexts for SOP Classes of the MWL Service Classes, and will respond to query requests from MWL-SCU.

### PRINT-SCU

PRINT-SCU is activated through the user interface when a user selects the currently displayed instances, and requests that they be sent to remote DICOM printer.

## Sequencing of Real-World Activities

All SCP activities are performed asynchronously in the background and not dependent on any sequencing.

SCU activities may be initiated or scheduled from the user interface, but are also performed asynchronously in the background. Other SCU activities may be triggered using routing rules or prior prefetching rules that occur in the background.

## RamSoft PACS AE Specification

The DICOM standard application context name and Implementation Identification Information are following:

Table 4.2

DICOM Application Context Name, Implementation Class and Version

| Application context name | 1.2.840.10008.3.1.1.1 |
| --- | --- |
| Implementation Class UID | 1.2.124.113540.1.6.0 |
| Implementation Version Name | RAMSOFT 6.0 |

# ECHO-SCP

### SOP CLASSES

ECHO-SCP provides Standard Conformance to the following SOP class:

Table 4.2.1.2

SOP Classes Supported by ECHO-SCP

| SOP Class Name | SOP Class UID |
| --- | --- |
| Verification SOP Class | 1.2.840.10008.1.1 |

#### Association Policies

        - General

ECHO-SCP accepts but never initiates associations.

#### Table 4.2.1.2.1 MAXIMUM PDU SIZE FOR ECHO-SCP

Configurable Parameter limited to 128 KB

Maximum PDU size

        - Number of Associations

ECHO-SCP will initiate at least one association with each destination for each task.

Each time the AE accepts an association, an available thread is assigned to complete the verification request. The maximum number of concurrent associations supported by RamSoft PACS is set by configuration. Once this number is reached, no more associations will be accepted until one of the threads become available.

#### Table 4.2.1.2.2

NUMBER OF ASSOCIATIONS FOR ECHO-SCP

Configurable Parameter

Maximum number of simultaneous associations

        - Asynchronous Nature

RamSoft PACS AE's do not provide asynchronous behavior. All association requests must be completed and acknowledged before a new operation can be performed.

#### Association Initiation Policy

ECHO-SCP does not initiate associations.

#### Association Acceptance Policy

When ECHO-SCP accepts an association, it will respond to echo requests. If the Called AE Title does not match the pre-configured AE Title in Station List and the option “Allow receiving without verifying AE Title credentials” is unchecked (See Paragraph 4.4.2Error! Reference source not found.), the association will be rejected.

        - Activity – Receive Echo Request

          - Description and Sequencing of Activities

As requests are received, they are responded to immediately.

          - Accepted Presentation Contexts

#### Table 4.2.1.4.1.2

Acceptable Presentation Contexts for ECHO-SCP and Receive ECHO Request

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| Verification | 1.2.840.10008.1.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP | None |

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance

            - SOP Specific Conformance to Verification SOP Class

ECHO-SCP provides standard conformance to the Verification Service Class

            - Presentation Context Acceptance Criterion

ECHO-SCP will always accept any Presentation Context for the supported SOP Classes with the supported Transfer Syntaxes. More than one proposed Presentation Context will be accepted for the same Abstract Syntax if the Transfer Syntax is supported, whether or not it is the same as another Presentation Context

            - Transfer Syntax Selection Policies

ECHO-SCP will select the first Transfer Syntax proposed by the client that is supported by the SCP, per Presentation Context.

# ECHO-SCU

#### SOP Classes

ECHO-SCU provides Standard Conformance to the following SOP Class:

Table 4.2.2.1

### SOP CLASSES SUPPORTED BY ECHO-SCU

| SOP Class Name | SOP Class UID |
| --- | --- |
| Verification SOP Class | 1.2.840.10008.1.1 |

#### Association Policies

        - General

ECHO-SCU initiates but never accepts associations.

#### Table 4.2.2.2.1 MAXIMUM PDU SIZE FOR ECHO-SCU

| Maximum PDU size | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

Table 4.2.2.2.2

### NUMBER OF ASSOCIATIONS FOR ECHO-SCU

| Maximum number of simultaneous associations | Unlimited |
| --- | --- |

        - Asynchronous Nature

RamSoft PACS AE's do not provide asynchronous behavior. All association requests must be completed and acknowledged before a new operation can be performed.

#### Association Initiation Policy

ECHO-SCU attempts to initiate a new association when the user clicks on Test Connection from the Station List.

        - Activity-Send Eco Request

          - Description and Sequencing of Activities

A single attempt will be made to verify DICOM connection to the remote AE.

          - Proposed Presentation Contexts

#### Table 4.2.2.3.1.2

PROPOSED PRESENTATION CONTEXTS FOR ECHO-SCU AND RECEIVE ECHO REQUEST

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| Verification | 1.2.840.10008.1.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCU | None |

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance

            - SOP Specific Conformance to verification SOP Class

ECHO-SCU provides standard conformance to the Verification Service Class.

            - Presentation Context Acceptance Criterion

ECHO-SCU does not accept associations.

            - Transfer Syntax Selection Policies

ECHO-SCU prefers Transfer Syntaxes in the order shown on the table 4.2.2.3.1.2 above.

#### Association Acceptance Policy

ECHO-SCU does not accept associations.

# STORAGE-SCP

#### SOP Classes

STORAGE-SCP provides Standard Conformance to the following SOP classes:

Table 4.2.3.1

### SOP CLASSES SUPPORTED BY STORAGE-SCP AND STORAGE-SCU

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Digital X-Ray Image Storage for Presentation | 1.2.840.10008.5.1.4.1.1.1.1 | Yes | Yes |
| Digital X-Ray Image Storage for Processing | 1.2.840.10008.5.1.4.1.1.1.1.1 | Yes | Yes |
| Computed Radiography Image Storage | 1.2.840.10008.5.1.4.1.1.1 | Yes | Yes |
| Digital Mammography X-Ray Image Storage for Presentation | 1.2.840.10008.5.1.4.1.1.1.2 | Yes | Yes |
| Digital Mammography X-Ray Image Storage for Processing | 1.2.840.10008.5.1.4.1.1.1.2.1 | Yes | Yes |
| Digital Intra Oral X-Ray Image Storage for Presentation | 1.2.840.10008.5.1.4.1.1.1.3 | Yes | Yes |
| Digital Intra Oral X-Ray Image Storage for Processing | 1.2.840.10008.5.1.4.1.1.1.3.1 | Yes | Yes |
| CT Image Storage | 1.2.840.10008.5.1.4.1.1.2 | Yes | Yes |
| RETIRED Ultrasound Multi Frame Image Storage | 1.2.840.10008.5.1.4.1.1.3 | Yes | Yes |
| Ultrasound Multi-frame Image Storage | 1.2.840.10008.5.1.4.1.1.3.1 | Yes | Yes |
| MR Image Storage | 1.2.840.10008.5.1.4.1.1.4 | Yes | Yes |
| Enhanced MR Image Storage | 1.2.840.10008.5.1.4.1.1.4.1 | Yes | Yes |
| MR Spectroscopy Image Storage | 1.2.840.10008.5.1.4.1.1.4.2 | Yes | Yes |
| RETIRED Nuclear Medicine Image Storage | 1.2.840.10008.5.1.4.1.1.5 | Yes | Yes |
| RETIRED Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6 | Yes | Yes |
| Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 | Yes | Yes |
| Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7 | Yes | Yes |
| Multi Frame Single Bit Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7.1 | Yes | Yes |

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Multi Frame Grayscale Byte Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7.2 | Yes | Yes |
| Multi Frame Grayscale Word Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7.3 | Yes | Yes |
| Multi Frame True Color Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7.4 | Yes | Yes |
| Basic Voice Audio Waveform Storage | 1.2.840.10008.5.1.4.1.1.9.4.1 | Yes | Yes |
| Hardcopy Grayscale Image Storage | 1.2.840.10008.5.1.1.29 | Yes | Yes |
| Hardcopy Color Image Storage | 1.2.840.10008.5.1.1.30 | Yes | Yes |
| Grayscale Softcopy Presentation State Storage | 1.2.840.10008.5.1.4.1.1.11.1 | Yes | Yes |
| X-Ray Angiographic Image Storage | 1.2.840.10008.5.1.4.1.1.12.1 | Yes | Yes |
| Enhanced XA Image Storage | 1.2.840.10008.5.1.4.1.1.12.1.1 | Yes | Yes |
| X-Ray Fluoroscopy Image Storage | 1.2.840.10008.5.1.4.1.1.12.2 | Yes | Yes |
| Enhanced XRF Image Storage | 1.2.840.10008.5.1.4.1.1.12.2.1 | Yes | Yes |
| RETIRED X-Ray Angiographic Bi-Plane Image Storage | 1.2.840.10008.5.1.4.1.1.12.3 | Yes | Yes |
| Nuclear Medicine Image Storage | 1.2.840.10008.5.1.4.1.1.20 | Yes | Yes |
| RETIRED VL Image Storage | 1.2.840.10008.5.1.4.1.1.77.1 | Yes | Yes |
| VL Endoscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.1 | Yes | Yes |
| VL Microscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.2 | Yes | Yes |
| VL Photographic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.4 | Yes | Yes |
| VL Slide Coordinates Microscopic Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.3 | Yes | Yes |
| RETIRED VL Multi Frame Image Storage | 1.2.840.10008.5.1.4.1.1.77.2 | Yes | Yes |
| Ophthalmic Photography 16 Bit Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.5.2 | Yes | Yes |
| Ophthalmic Photography 8 Bit Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.5.1 | Yes | Yes |
| Ophthalmic Tomography Image Storage | 1.2.840.10008.5.1.4.1.1.77.1.5.4 | Yes | Yes |
| Enhanced XRF Image Storage | 1.2.840.10008.5.1.4.1.1.12.2.1 | Yes | Yes |
| Encapsulated PDF Storage | 1.2.840.10008.5.1.4.1.1.104.1 | Yes | Yes |
| Enhanced CT Image Storage | 1.2.840.10008.5.1.4.1.1.2.1 | Yes | Yes |
| Basic Text SR | 1.2.840.10008.5.1.4.1.1.88.11 | Yes | Yes |
| Enhanced SR | 1.2.840.10008.5.1.4.1.1.88.22 | Yes | Yes |
| Comprehensive SR | 1.2.840.10008.5.1.4.1.1.88.33 | Yes | Yes |

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Mammography CAD SR | 1.2.840.10008.5.1.4.1.1.88.50 | Yes | Yes |
| X-Ray Radiation Dose SR Storage | 1.2.840.10008.5.1.4.1.1.88.67 | Yes | Yes |
| Raw Data Storage | 1.2.840.10008.5.1.4.1.1.66 | Yes | Yes |
| Key Object Selection | 1.2.840.10008.5.1.4.1.1.88.59 | Yes | Yes |
| PET Image Storage | 1.2.840.10008.5.1.4.1.1.128 | Yes | Yes |
| Enhanced PET Image Storage | 1.2.840.10008.5.1.4.1.1.130 | Yes | Yes |
| Breast Tomosynthesis Image Storage | 1.2.840.10008.5.1.4.1.1.13.1.3 | Yes | Yes |
| Fuji Computed Radiography Image Storage | 1.2.392.200036.9125.1.1.2 | Yes | Yes |
| Media Storage Directory Storage | 1.2.840.10008.1.3.10 | Yes | Yes |

No changes are required to the Storage SOP Class transfer syntax tables.

Lossless and lossy image compression (JPEG, JPEG-LS, JPEG 2000, etc.) continue to be negotiated per Storage SOP Class as documented elsewhere in the conformance statement.

#### Association Policies

        - General

STORAGE-SCP accepts but never initiates associations.

Table 4.2.3.2.1

### MAXIMUM PDU SIZE FOR STORAGE-SCP

| Maximum PDU size | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

Each time the AE accepts an association, an available thread is assigned to complete the transfer of medical images. The maximum number of concurrent associations supported by RamSoft PACS is set by configuration. Once this number is reached, no more associations will be accepted until one of the threads become available.

#### Table 4.2.3.2.2

NUMBER OF ASSOCIATIONS FOR STOAGE-SCP

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

        - Asynchronous Nature

PowerServer AE's do not provide asynchronous behavior. All association requests must be completed and acknowledged before a new operation can be performed.

#### Association Initiation Policy

STORAGE-SCP does not initiate associations.

#### Association Acceptance Policy

An association will be accepted with an external SCU if the requesting SCU provides valid parameters. Valid parameters include a valid presentation context. The AE's title is verified with RamSoft PACS’ Station List, unless promiscuous receiving is enabled.

        - Activity – Receive Storage Request

          - Description and Sequencing of Activities

As requests are received, they are responded to immediately.

          - Accepted Presentation Contexts

Table 4.2.3.4.1.2-1

### PRESENTATION CONTEXTS FOR STORAGE-SCP AND STORAGE-SCU

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.3.1 | See Table 4.2.3.1 | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP, SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP, SCU | None |
|  |  | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP, SCU | None |
|  |  | Explicit VR Big Endian | 1.2.840.10008.1.2.2 | SCP, SCU | None |
|  |  | JPEG Lossy Baseline 8-bit Image Compression | 1.2.840.10008.1.2.4.50 | SCP, SCU | None |
|  |  | JPEG Lossy Extended 12-bit Image Compression | 1.2.840.10008.1.2.4.51 | SCP, SCU | None |

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
|  |  | JPEG Lossless Non-hierarchical (Process 14) | 1.2.840.10008.1.2.4.57 | SCP, SCU | None |
|  |  | JPEG Lossless, Non-hierarchical First-Order Prediction (Process 14) | 1.2.840.10008.1.2.4.70 | SCP, SCU | None |
|  |  | JPEG-LS Lossless Image Compression | 1.2.840.10008.1.2.4.80 | SCP, SCU | None |
|  |  | JPEG 2000 Lossless Only | 1.2.840.10008.1.2.4.90 | SCP, SCU | None |
|  |  | JPEG 2000 | 1.2.840.10008.1.2.4.91 | SCP, SCU | None |
|  |  | RLE Lossless Transfer Syntax | 1.2.840.10008.1.2.5 | SCP, SCU | None |

Note: All Storage SOP classes in Table 4.2.3.1 support all of the above transfer syntaxes except for the SOP Classes listed in Table 4.2.3.4.1.2-2

#### Table 4.2.3.4.1.2-2

The list of SOP classes that support only the first three entries from Table: 4.2.3.4.1.2-1

| SOP Class Name | SOP Class UID |
| --- | --- |
| Multiframe Single Bit Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7.1 |
| Basic Voice Audio Waveform Storage | 1.2.840.10008.5.1.4.1.1.9.4.1 |
| Grayscale Softcopy Presentation State Storage | 1.2.840.10008.5.1.4.1.1.11.1 |
| Basic Text SR | 1.2.840.10008.5.1.4.1.1.88.11 |

| Enhanced SR | 1.2.840.10008.5.1.4.1.1.88.22 |
| --- | --- |
| Comprehensive SR | 1.2.840.10008.5.1.4.1.1.88.33 |
| Mammography CAD SR | 1.2.840.10008.5.1.4.1.1.88.50 |
| X-Ray Radiation Dose SR Storage | 1.2.840.10008.5.1.4.1.1.88.67 |
| Key Object Selection | 1.2.840.10008.5.1.4.1.1.88.59 |
| Encapsulated PDF Storage | 1.2.840.10008.5.1.4.1.1.104.1 |

            - Extended Negotiation

No extended negotiation is performed, though STORAGE-SCP:

- is a Level 2 Storage SCP (Full – does not discard any data elements)

- received data elements may be coerced through configurable parameters, Field Mapping and DICOM Scripting, all of which are configured on the Station List

          - SOP Specific Conformance

            - SOP Specific Conformance to STORAGE SOP Classes

STORAGE-SCP provides standard conformance to the Storage Service Classes

              - SOP Specific Conformance – Grayscale Softcopy Presentation State Storage

Presentation related attributes are derived from the displayed image and included in the IOD when image parameters are saved as selected by the user or when the displayed study is closed, if automatic saving of presentation states is enabled for the user. All Image Storage SOP Classes listed in Table 4.2.3.1 may be referenced by instances of the Grayscale Softcopy Presentation State Storage SOP Class.

              - SOP Specific Conformance – Structured Reporting Storage

No images or other composite object Storage SOP Classes may be referenced by newly created instances of Structured Reporting Storage SOP Class. Only Basic Text Structured Reporting Storage SOP Classes are created. A new SOP Instance UID is created for Structured Reporting Storage SOP Classes whenever any attribute of the SOP Class is modified.

Structured Reporting related attributes are rendered by the SCP as a PDF document using a configurable template. Image or other object storage SOP instances referenced by instances of the Structured Reporting SOP Class are ignored and not used to influence the display of the image or object. The SCP never renders content items with Rendering Intent Concept modifier set to “Presentation Optional".

              - SOP Specific Conformance – Media Storage Directory Storage

RamSoft DICOM AE supports the Media Storage Application Profiles shown in the Table 1-2.

STORAGE-SCP provides standard conformance to the Interchange option of the Media Storage Service Class with Directory Information as a File-set Creator (FSC), File-set Updater (FSU). The directory information is present. No optional standard keys are included in directory records. All Storage SOP classes listed in

Table 4.2.3.1 can also functioning as Media Storage SOP classes.

            - Presentation Context Acceptance Criterion

STORAGE-SCP will always accept any Presentation Context for the supported SOP Classes with the supported Transfer Syntaxes. More than one proposed Presentation Context will be accepted for the same Abstract Syntax if the Transfer Syntax is supported, whether or not it is the same as another Presentation Context

            - Transfer Syntax Selection Policies

PowerServer transfer syntax selection policy can be configured in the Station List. The support for any of Uncompressed, Lossy Compressed and Lossless Compressed transfer syntaxes may be enabled or disabled. The preference of each type of compression may also be set in the Station List.

            - Response Status

STORAGE-SCP will behave as described in the Table 4.2.3.4.1.3.4 below when generating the C-STORE response command message.

#### Table 4.2.3.4.1.3.4

RESPONSE STATUS FOR STORAGE-SCP AND RECEIVE STORAGE REQUEST

| Service Status | Further Meaning | Status Codes | Reason |
| --- | --- | --- | --- |
| Refused | Out of Resources | A700 | Not enough disk space. Image will not be saved. Association will be aborted. |
| Error | SOP Class Not Supported | A800 | Requested in Association SOP class is not supported. Association will be aborted. |
|  | Cannot understand | C000 | The Storage AE couldn’t parse the data set into elements. Association will be aborted |
| Success | Success | 0000 |  |

# STORAGE-SCU

### SOP CLASSES

STORAGE-SCU provides Standard Conformance to the list of SOP Classes in Table 4.2.3.1

#### Association Policies

        - General

STORAGE-SCU initiates but never accepts associations. Maximum PDU Size received as a SCP for STORAGE-SCU is identical to STORAGE-SCP and shown in Table 4.2.3.2.1

No changes are required to the Storage SOP Class transfer syntax tables.

Lossless and lossy image compression (JPEG, JPEG-LS, JPEG 2000, etc.) continue to be negotiated per Storage SOP Class as documented elsewhere in the conformance statement.

        - Number of Associations

#### Table 4.2.4.2.2

NUMBER OF ASSOCIATIONS FOR STORAGE-SCU

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

        - Asynchronous Nature

PowerServer AE's do not provide asynchronous behavior. All association requests must be completed and acknowledged before a new operation can be performed.

#### Association Initiation Policy

STORAGE-SCU will initiate a new association whenever the user requests or sends a study. Multiple studies to a single destination may be transmitted in a single association. In Station List, a preference for compressed transfer syntaxes can be configured for each station. Default compression can be set through Settings/Server Settings/Compression. The order in which compressed transfer syntaxes are to be presented is also configured through Station List. Lossy compressed images will be sent as is, so long as the receiving station supports that transfer syntax. Otherwise, images will be decoded and encoded as necessary.

An association will be accepted with an external SCU if the requesting SCU provides valid parameters. Valid parameters include a valid presentation context. The AE's title is verified with RamSoft PACS’ Station List unless promiscuous receiving is enabled

        - Activity – Send Storage Request

          - Description and Sequencing of Activities

The AE will initiate a new association whenever the user requests or sends a study. Multiple studies to a single destination may be transmitted in a single association. In Station List, a preference for compressed transfer syntaxes can be configured for each station. Default compression can be set through Settings/Server Settings/Compression. The order in which compressed transfer syntaxes are to be presented is also configured through Station List. Lossy compressed images will be sent as is, so long as the receiving station supports that transfer syntax. Otherwise, images will be decoded and encoded as necessary.

          - Proposed Presentation Contexts

Proposed Presentation Contexts are illustrated in Table 4.2.3.4.1.2-1 and Table 4.2.3.4.1.2-2

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance

            - SOP Specific Conformance to STORAGE SOP Classes

STORAGE-SCU provides standard conformance to the Storage Service Classes.

After a successful C-STORE response from the SCP, the AE will continue to send any unsent images or SR documents belonging to the same study. If a particular image cannot be sent in the current transfer syntax (due to limitations of some compressed transfer syntaxes), the association will be released and a new association will be created. If a particular image and the association have the same transfer syntax, then the image is sent keeping the original encapsulation format (compressed/uncompressed) even if the compression is disabled for association (as a user data configuration). If an unsuccessful C-STORE response is received from the SCP, the AE will record the failure and close the association. The study will remain in the Transmit List for retry at the next available time. Warnings in the C-STORE response from the SCP are ignored.

If PowerServer originally acquired or created the image, all mandatory modules of the image storage objects are provided. All type 1 and type 2 data types are provided. Optional data types may not be provided. If the image was originally acquired through a DICOM device, RamSoft PACS saves all tags received with the image. Thus, when these images are transmitted, all optional tags originally received by RamSoft PACS will also be transmitted. If the acquired image has no DICOM SOP class item, then RamSoft PACS stores the images as a Secondary Capture Image Storage object.

              - SOP Specific Conformance – Grayscale Softcopy Presentation State Storage

This AE provides standard conformance to Grayscale Softcopy Presentation State Storage. RamSoft PowerServer makes all mandatory presentation attributes available for application to the referenced images at the discretion of the user for all Image Storage SOP Classes listed in Table 4.2.3.1

              - SOP Specific Conformance – Structured Report Storage

This AE provides standard conformance to Structured Reporting Storage. RamSoft renders the structured report related attributes as a PDF document with a configurable template. No image or other composite object Storage SOP Classes will be displayed or otherwise rendered.

Text report body is stored on the first text node on the document tree. RamSoft PowerServer stores binary data on a private tag (3113, 1010). The binary document size and type are specified on private tags (3113, 1020) and (3111, 1010). When a structured report has an associated Microsoft Word document, file data and size of the template file are stored on private tags (3113, 1030) and (3113, 1040).

              - SOP Specific Conformance – Media Storage Directory Storage

STORAGE-SCU provides standard conformance to the Interchange option of the Media Storage Service Class with Directory Information as a File-set Reader (FSR). No optional standard keys are included in directory records. All Storage SOP classes listed in Table 4.2.3.1 can also function as Media Storage SOP classes. RamSoft PACS conforms to the application profiles listed in Table 4.2.3.4.1.3.1.3

            - Presentation Context Acceptance Criterion

STORAGE-SCU does not accept associations.

            - Transfer Syntax Selection Policies

PowerServer transfer syntax selection policy can be configured in the Station List (Section 4.4.2Error! Reference source not found.). The support for any of Uncompressed, Lossy Compressed and Lossless Compressed transfer syntaxes may be enabled or disabled. The preference of each type of compression may also be set in the “Station List”.

            - Response Status

STORAGE-SCU will behave as described in the Table below in response to the status returned in the C-STORE response command message.

#### Table 4.2.3.4.1.3.4

RESPONSE STATUS FOR STORAGE-SCU AND RECEIVE STORAGE REQUEST

| Service Status | Further Meaning | Status Codes | Behavior |
| --- | --- | --- | --- |
| Refused | Out of Resources | A700 | Abort Association |
| Error | SOP Class Not Supported | A800 | Abort Association |
|  | Data Set does not match SOP Class | A900 | Abort Association |
|  | Cannot understand | C000 | Abort Association |
| Warning | Coercion of Data Elements | B000 | Ignored |
|  | Data Set does not match SOP Class | B007 | Ignored |
|  | Elements Discarded | B006 | Ignored |

| Service Status | Further Meaning | Status Codes | Behavior |
| --- | --- | --- | --- |
| Success | Success | 0000 |  |

#### Association Acceptance Policy

STORAGE-SCU does not accept associations.

# STORAGE-SCU-COMMITMENT

#### SOP Classes

STORAGE-SCU-COMMITMENT provides Standard Conformance to the following SOP Class:

#### Table 4.2.5.1

SOP CLASSES SUPPORTED BY STORAGE-SCU-COMMITMENT

| SOP Class Name | SOP Class UID |
| --- | --- |
| Storage Commitment Push Model | 1.2.840.10008.1.20.1 |

        - Proposed Presentation Contexts

#### Table 4.2.5.1.1

PRESENTATION CONTEXTS FOR STORAGE-SCU-COMMITMENT

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| Storage Commitment Push Model | 1.2.840.10008.1.20.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCU | None |

#### Association Initiation Policy

        - Activity-Send Images

          - SOP Specific Conformance – Storage Commitment Push Model

STORAGE-SCU-COMMITMENT AE provides standard conformance to this SOP Class as an SCU. After transmission complete, N-ACTION is issued with list of objects that have been transmitted, when peer AE is configured to support Storage Commitment.

If N-EVENT-REPORT response is received on the same association, processing is completed and Receiving of If N-EVENT-REPORT is committed. Otherwise, the STORAGE-SCU-COMMITMENT will wait until N-EVENT-REPORT request is received in another association.

# STORAGE-SCP-COMMITMENT

#### SOP Classes

STORAGE-SCP-COMMITMENT provides Standard Conformance to the following SOP Class:

#### Table 4.2.6.1

SOP CLASSES SUPPORTED BY STORAGE-SCP-COMMITMENT

| SOP Class Name | SOP Class UID |
| --- | --- |
| Storage Commitment Push Model | 1.2.840.10008.1.20.1 |

        - Proposed Presentation Contexts

#### Table 4.2.6.1.1

PRESENTATION CONTEXTS FOR STORAGE-SCP-COMMITMENT

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| Storage Commitment Push Model | 1.2.840.10008.1.20.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP | None |

#### Association Initiation Policy

        - Activity-Receiving Images

          - SOP Specific Conformance – Storage Commitment Push Model

STORAGE-SCP-COMMITMENT AE provides standard conformance to this SOP Class as an SCP. After receiving N-ACTION request N-EVENT-REPORT response is sent to STORAGE-SCU-COMMITMENT. This response can sent in some association as association initiated be STORAGE-SCU-COMMITMENT or in another association. The setting “Send Event Report Message on the same Association” (Section 4.4.2 – Figure 4.4.2-5 Station List  – Receive Service ) manages this behavior.

# FIND-SCP

#### SOP Classes

FIND-SCP provides Standard Conformance to the following SOP Classes:

#### Table 4.2.7.1

SOP CLASSES SUPPORTED BY FIND-SCP AND FIND-SCU

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Study Root Query/Retrieve Information Model – FIND | 1.2.840.10008.5.1.4.1.2.2.1 | Yes | Yes |
| Patient Root Query/Retrieve Information Model – FIND | 1.2.840.10008.5.1.4.1.2.1.1 | No | Yes |
| Patient/Study Only Query/Retrieve Information Model - FIND | 1.2.840.10008.5.1.4.1.2.3.1 | No | Yes |

### Transfer Syntax Support

#### For DIMSE Query/Retrieve operations (C-FIND), Omega AI supports the following transfer syntaxes for command and identifier datasets:

#### Table 4.2.7.2‑1

#### Supported Transfer Syntaxes for FIND-SCP

| #### Transfer Syntax | #### UID | #### Supported |
| --- | --- | --- |
| #### Implicit VR Little Endian | #### 1.2.840.10008.1.2 | #### Yes |
| #### Explicit VR Little Endian | #### 1.2.840.10008.1.2.1 | #### Yes |
| #### Deflated Explicit VR Little Endian | #### 1.2.840.10008.1.2.1.99 | #### Not proposed |
| #### Explicit VR Big Endian | #### 1.2.840.10008.1.2.2 | #### Not proposed |

#### Rationale: Although Deflated Explicit VR Little Endian is a valid DICOM transfer syntax, it is optional and has been observed to cause interoperability issues in DIMSE Query/Retrieve workflows, particularly with multi-response operations such as C-FIND and C-MOVE. As these operations exchange small command datasets where compression provides no practical benefit, Omega AI restricts DIMSE Query/Retrieve negotiation to uncompressed Little Endian transfer syntaxes.

#### Association Policies

        - General

FIND-SCP accepts but never initiates associations.

#### Table 4.2.7.3.1

MAXIMUM PDU SIZE RECEIVED AS A SCP FOR FIND-SCP

Configurable Parameter limited to 128 KB

Maximum PDU size received

        - Number of Associations

Each time the FIND-SCP accepts an association, an available thread is assigned to complete the find (query) request. The maximum number of concurrent associations supported by RamSoft PACS is set by configuration. Once this number is reached, no more associations will be accepted until one of the threads become available.

#### Table 4.2.7.3.2

NUMBER OF ASSOCIATIONS AS A SCP FOR FIND-SCP

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

#### Association Initiation Policy

FIND-SCP does not initiate associations.

#### Association Acceptance Policy

When FIND-SCP accepts an association, it will respond to query requests. If the Called AE Title does not match the pre-configured AE Title in Station List and option “Allow receiving without verifying AE Title credentials” (See Paragraph 4.4.2Error! Reference source not found.) is unchecked the association will be rejected.

        - Activity – Receive Query Request

          - Description and Sequencing of Activities

As requests are received, the Database is queried for result set.

          - Accepted Presentation Contexts

Table 4.2.7.5.1.2

### ACCEPTABLE AND PROPOSED PRESENTATION CONTEXTS FOR FIND-SCP AND FIND-SCU

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.7.1 | See Table 4.2.7.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP,SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP,SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP,SCU | None |

            - Extended Negotiation

No extended negotiation is performed. In particular, relational queries are not supported.

          - SOP Specific Conformance

            - SOP Specific Conformance to C-FIND SOP Classes

FIND-SCP provides standard conformance to the supported C-FIND SOP Classes. The maximum number of records returned is limited to 500.

              - The types of Matching supported by the FIND-SCP and FIND-SCU

An "S" indicates the identifier attribute uses Single Value Matching, an “R” indicates Range Matching, an “*” indicates wildcard matching, a ‘U’ indicates Universal Matching, and an ‘L’ indicates that UID lists are sent. “NONE” indicates that no matching is supported, but that values for this Element are requested to be returned (i.e. universal matching),

and “UNIQUE” indicates that this is the Unique Key for that query level, in which case Universal Matching or Single

Value Matching is used depending on the query level.

              - Option for return requested attributes

Setting “Return C-FIND Request Attributes Only” (Section 4.4.2 – Error! Reference source not found. Station List – Receive Service) gives the possibility to return only requested attributes (“Checkbox selected”) or all attributes (“Checkbox unselected”). This setting is related to all classes shown in Table 4.2.7.1

              - Patient Root Query Information Model

The fixed set of attributes is returned in a C-FIND response as illustrated in Table 4.2.7.5.1.3.1.3

#### Table 4.2.7.5.1.3.1.3

PATIENT ROOT REQUEST IDENTIFIER FOR FIND-SCP

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| PATIENT Level |  |  |  |
| Patient’s ID | (0010,0020) | S,*,U | YES |
| Patient’s Issuer | (0010,0021) | S,*,U | YES |
| Patient’s Name | (0010,0010) | S,*,U | YES |
| Patient's Birth Date | (0010,0030) | S,*,U,R | YES |
| Patient's Sex | (0010,0040) | S,*,U | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Patient’s Insurance Plan Code Sequence | (0010,0050) | NONE | YES |
| >Code Value | (0008,0100) | NONE | YES |
| >Patient’s Primary Language Code Sequence | (0010,0101) | NONE | YES |
| >> Code Value | (0008,0100) | NONE | YES |
| STUDY Level |  |  |  |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Patient’s ID | (0010,0020) | S,*,U | YES |
| Patient’s Issuer | (0010,0021) | S,*,U | YES |
| Patient’s Name | (0010,0010) | S,*,U | YES |
| Patient's Birth Date | (0010,0030) | S,*,U,R | YES |
| Patient's Sex | (0010,0040) | S,*,U | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Patient’s Insurance Plan Code Sequence | (0010,0050) | NONE | YES |
| >Code Value | (0008,0100) | NONE | YES |
| >Patient’s Primary Language Code Sequence | (0010,0101) | NONE | YES |
| >> Code Value | (0008,0100) | NONE | YES |
| Modalities In Study | (0008,0061) | S,*,U,R | YES |
| Manufacturer | (0008,0070) | NONE | YES |
| Institution’s Name | (0008,0080) | S,*,U | YES |
| Referring Physician’s Name | (0008,0090) | NONE | YES |
| Manufacturer Model Name | (0008,1090) | NONE | YES |
| Software Versions | (0018,1020) | NONE | YES |
| Study ID | (0020,0010) | S,*,U,R | YES |
| Study Description | (0008,1030) | NONE | YES |
| Modalities In Study | (0008,0061) | S,*,U | YES |
| Time Zone Offset From UTC | (0008,0201) | NONE | YES |
| Study Date | (0008,0020) | S,*,U,R | YES |
| Study Time | (0008,0030) | S,*,U,R | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Accession Number | (0008,0050) | S,*,U | YES |
| Number Of Study Related Series | (0020,1206) | NONE | YES |
| Number Of Study Related Instances | (0020,1208) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | S,U | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| Requesting Physician | (0032,1032) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| StudyInstanceUID | (0020,000D) | UNIQUE | Study Instance UID |
| SERIES Level |  |  |  |
| Patient’s ID | (0010,0020) | NONE | YES |
| Patient’s Issuer | (0010,0021) | NONE | YES |
| Patient’s Name | (0010,0010) | NONE | YES |
| Patient's Birth Date | (0010,0030) | NONE | YES |
| Patient's Sex | (0010,0040) | NONE | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Study ID | (0020,0010) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | NONE | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| Study Date | (0008,0020) | NONE | YES |
| Study Time | (0008,0030) | NONE | YES |
| Accession Number | (0008,0050) | NONE | YES |
| Number Of Series Related Instances | (0020,1209) | NONE | YES |
| StudyInstanceUID | (0020,000D) | UNIQUE | YES |
| SeriesInstanceUID | (0020,000E) | UNIQUE | YES |
| Body Part Examined | (0018,0015) | NONE | YES |
| Series Number | (0020,0011) | S,*,U | YES |
| Laterality | (0020,0060) | NONE | YES |
| Modality | (0008,0060) | NONE | YES |
| IMAGE Level |  |  |  |
| Patient’s ID | (0010,0020) | NONE | YES |
| Patient’s Issuer | (0010,0021) | NONE | YES |
| Patient’s Name | (0010,0010) | NONE | YES |
| Patient's Birth Date | (0010,0030) | NONE | YES |
| Patient's Sex | (0010,0040) | NONE | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Study ID | (0020,0010) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | NONE | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| Study Date | (0008,0020) | NONE | YES |
| Study Time | (0008,0030) | NONE | YES |
| Accession Number | (0008,0050) | NONE | YES |
| Body Part Examined | (0018,0015) | NONE | YES |
| Series Number | (0020,0011) | NONE | YES |
| Laterality | (0020,0060) | NONE | YES |
| Modality | (0008,0060) | NONE | YES |
| SOPClassUID | (0008,0016) | NONE | YES |
| SOPInstance UID | (0008,0018) | UNIQUE | YES |
| Instance Number | (0020,0013) | S,*,U | YES |
| Concept Name Code Sequence | (0040, A043) | NONE | YES |
| >Code Value | (0008,0100) | NONE | YES |
| >Coding Scheme Designator | (0008,0102) | NONE | YES |
| >Code Meaning | (0008,0104) | NONE | YES |
| Common to all query levels |  |  |  |
| Retrieve AETitle | (0008,0054) | NONE | YES |

              - Study Root Query Information Model

The attributes that are requested and additional attributes are returned in a C-FIND response as illustrated in Table 0

Table 0

### STUDY ROOT REQUEST IDENTIFIER FOR FIND-SCP

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| STUDY Level |  |  |  |
| Patient’s ID | (0010,0020) | S,*,U | YES |
| Patient’s Issuer | (0010,0021) | S,*,U | YES |
| Patient’s Name | (0010,0010) | S,*,U | YES |
| Patient's Birth Date | (0010,0030) | S,*,U,R | YES |
| Patient's Sex | (0010,0040) | S,*,U | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Patient’s Insurance Plan Code Sequence | (0010,0050) | NONE | YES |
| >Code Value | (0008,0100) | NONE | YES |
| >Patient’s Primary Language Code Sequence | (0010,0101) | NONE | YES |
| >> Code Value | (0008,0100) | NONE | YES |
| Modalities In Study | (0008,0061) | S,*,U | YES |
| Manufacturer | (0008,0070) | NONE | YES |
| Institution’s Name | (0008,0080) | S,*,U | YES |
| Referring Physician’s Name | (0008,0090) | NONE | YES |
| Manufacturer Model Name | (0008,1090) | NONE | YES |
| Software Versions | (0018,1020) | NONE | YES |
| Study ID | (0020,0010) | S,*,U | YES |
| Study Description | (0008,1030) | NONE | YES |
| Time Zone Offset From UTC | (0008,0201) | NONE | YES |
| Study Date | (0008,0020) | S,*,U,R | YES |
| Study Time | (0008,0030) | S,*,U,R | YES |
| Accession Number | (0008,0050) | S,*,U | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Number Of Study Related Series | (0020,1206) | NONE | YES |
| Number Of Study Related Instances | (0020,1208) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | S,U | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| Requesting Physician | (0032,1032) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| StudyInstanceUID | (0020,000D) | UNIQUE | Study Instance UID |
| SERIES Level |  |  |  |
| Patient’s ID | (0010,0020) | NONE | YES |
| Patient’s Issuer | (0010,0021) | NONE | YES |
| Patient’s Name | (0010,0010) | NONE | YES |
| Patient's Birth Date | (0010,0030) | NONE | YES |
| Patient's Sex | (0010,0040) | NONE | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Study ID | (0020,0010) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | NONE | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| Study Date | (0008,0020) | NONE | YES |
| Study Time | (0008,0030) | NONE | YES |
| Accession Number | (0008,0050) | NONE | YES |
| Number Of Series Related Instances | (0020,1209) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| StudyInstanceUID | (0020,000D) | UNIQUE | YES |
| SeriesInstanceUID | (0020,000E) | UNIQUE | YES |
| Body Part Examined | (0018,0015) | NONE | YES |
| Series Number | (0020,0011) | S,*,U | YES |
| Laterality | (0020,0060) | NONE | YES |
| Modality | (0008,0060) | NONE | YES |
| IMAGE Level |  |  |  |
| Patient’s ID | (0010,0020) | NONE | YES |
| Patient’s Issuer | (0010,0021) | NONE | YES |
| Patient’s Name | (0010,0010) | NONE | YES |
| Patient's Birth Date | (0010,0030) | NONE | YES |
| Patient's Sex | (0010,0040) | NONE | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Address | (0010,1040) | NONE | YES |
| Country of Residence | (0010,2150) | NONE | YES |
| Patient’s Telephone Numbers | (0010,2154) | NONE | YES |
| Study ID | (0020,0010) | NONE | YES |
| RETIRED StudyStatusID | (0032,000A) | NONE | YES |
| RETIRED StudyPriorityID | (0032,000C) | NONE | YES |
| RETIRED Study Comments | (0032,4000) | NONE | YES |
| Route of AdmissionID | (0038,0016) | NONE | YES |
| Current Patient Location | (0038,0300) | NONE | YES |
| RETIRED Interpretation Transcriber | (4008,010A) | NONE | YES (Transcriptionist) |
| Study Date | (0008,0020) | NONE | YES |
| Study Time | (0008,0030) | NONE | YES |
| Accession Number | (0008,0050) | NONE | YES |
| Body Part Examined | (0018,0015) | NONE | YES |
| Series Number | (0020,0011) | NONE | YES |
| Laterality | (0020,0060) | NONE | YES |
| Modality | (0008,0060) | NONE | YES |
| SOPClassUID | (0008,0016) | NONE | YES |
| SOPInstance UID | (0008,0018) | UNIQUE | YES |
| Instance Number | (0020,0013) | S,*,U | YES |
| Concept Name Code Sequence | (0040, A043) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| >Code Value | (0008,0100) | NONE | YES |
| >Coding Scheme Designator | (0008,0102) | NONE | YES |
| >Code Meaning | (0008,0104) | NONE | YES |
| Common to all query levels |  |  |  |
| Retrieve AETitle | (0008,0054) | NONE | YES |

              - Patient / Study Only Query Information Model

This Query information level is identical to Patient Root Query Information Level in RamSoft PowerServer Application.

            - Presentation Context Acceptance Criterion

FIND-SCP will accept the Presentation Contexts given in Table 4.2.7.5.1.2

            - Transfer Syntax Selection Policies

FIND-SCP will prefer explicit transfer syntaxes over implicit transfer syntaxes. Duplicate Presentation Contexts will not be accepted. If it is offered three identical Presentation Contexts, each of which offers any of the three acceptable Transfer Syntaxes, it will accept all Presentation Contexts, but with different Transfer Syntaxes in each. Other transfer syntax policies may be configured in the Station List.

            - Response Status

Table 4.2.7.5.1.3.1.4

### RESPONSE STATUS FOR FIND-SCP AND RECEIVE QUERY REQUEST

| Service Status | Further Meaning | Status Codes | Reason |
| --- | --- | --- | --- |
| Failure | Refused: Out of resources | A700 | Not enough resources to complete query. Association will be dropped. |
|  | Refused: SOP Class requested in Association is not supported. | A800 | Requested SOP Class is not Supported |
|  | Identifier does not match SOP Class | A900 | Sent if an invalid query retrieve level is set |
|  | Unable to process | C000 | Sent if internal database query is Unsuccessful. |

| Service Status | Further Meaning | Status Codes | Reason |
| --- | --- | --- | --- |
| Cancel | Matching terminated due to Cancel | FE00 | Sent when a C-CANCEL Request is received |
| Pending | Matches are continuing - Current Match is supplied and any Optional Keys were supported in the same manner as Required keys | FF00 | Sent |
|  | Matches are continuing - warning that one or more Optional Keys were not supported for existence for this identifier | FF01 | Sent |
| Success | Matching is complete | 0000 | Sent |

# FIND-SCU

#### SOP Classes

FIND-SCU provides Standard Conformance to the SOP Classes shown in Table 4.2.7.1

### Transfer Syntax Support

#### For DIMSE C-FIND operations, Omega AI supports the following transfer syntaxes for command and identifier datasets:

Table 4.2.8.2-1

Supported Transfer Syntaxes for FIND-SCU

| Transfer Syntax | UID | Supported |
| --- | --- | --- |
| Implicit VR Little Endian | 1.2.840.10008.1.2 | Yes |
| Explicit VR Little Endian | 1.2.840.10008.1.2.1 | Yes |
| Deflated Explicit VR Little Endian | 1.2.840.10008.1.2.1.99 | Not proposed |
| Explicit VR Big Endian | 1.2.840.10008.1.2.2 | Not proposed |

Rationale: See Section 4.2.7.2 for rationale on transfer syntax restrictions applicable to all DIMSE Query/Retrieve operations.

#### Association Policies

        - General

FIND-SCU initiates but never accepts associations.

#### Table 4.2.8.3.1 MAXIMUM PDU SIZE FOR FIND-SCU

| Maximum PDU size | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

#### Table 4.2.8.3.2 MAXIMUM PDU SIZE FOR FIND-SCU

Configurable Parameter

Maximum number of simultaneous associations

        - Asynchronous Nature

FIND-SCU will only allow a single outstanding operation on an Association.

#### Association Acceptance Policy

FIND-SCU does not accept associations.

#### Association Initiation Policy

FIND-SCU attempts to initiate a new association when the user performs the query action from the user interface. (Advanced Search ->Remote Studies).

        - Activity – Query Remote AE

          - Description and Sequencing of Activities

The attempt will be made to query the remote AE.

          - Proposed Presentation Contexts

Proposed Presentation Contexts for FIND-SCU are identical to Accepted Presentation Contexts of FIND-SCP and illustrated in Table 4.2.7.5.1.2

            - Extended Negotiation

No extended negotiation is performed. In particular, relational queries are not supported

          - SOP Specific Conformance

            - SOP Specific Conformance to C-FIND SOP Class

FIND-SCU provides standard conformance to the supported C-FIND SOP Classes. Only Study Root Query Information Model is supported and is used at the STUDY level. The requested attributes are illustrated in Table 4.2.8.5.1.3.1

#### Table 4.2.8.5.1.3.1

STUDY ROOT REQUEST IDENTIFIER FOR FIND-SCU

| Name | Tag | Types of Matching |
| --- | --- | --- |
| STUDY Level |  |  |
| Patient’s ID | (0010,0020) | S,*,U |

| Patient’s Issuer | (0010,0021) | S,*,U |
| --- | --- | --- |
| Patient’s Name | (0010,0010) | S,*,U |
| Other Patient's ID's | (0010,1000) | S,*,U |
| Modalities In Study | (0008,0061) | S,*,U |
| Institution Name | (0008,0080) | U,* |
| Study Date | (0008,0020) | U,R |
| Study ID | (0020,0010) | S,*,U,R |
| Accession Number | (0008,0050) | U,* |
| RETIRED StudyStatusID | (0032,000A) | S,*,U |
| RETIRED StudyPriorityID | (0032,000C) | S,*,U |

            - Presentation Context Acceptance Criterion

FIND-SCU will accept the Presentation Contexts given in Table 4.2.7.5.1.2

            - Transfer Syntax Selection Policies

FIND-SCU will prefer explicit transfer syntaxes over implicit transfer syntaxes.

            - Response Status

FIND-SCU will behave as described in Table 4.2.8.5.1.3.4 in response to the status returned in the C-FIND response command messages.

Table 4.2.8.5.1.3.4

### RESPONSE STATUS FOR FIND-SCU AND QUERY REMOTE AE REQUEST

| Service Status | Further Meaning | Status Codes | Reason |
| --- | --- | --- | --- |
| Failure | Refused: Out of resources | A700 | Not enough resources to complete query. Association will be dropped. |
|  | Refused: SOP Class requested in Association is not supported. | A800 | Requested SOP Class is not Supported Association will be dropped. |
|  | Identifier does not match SOP Class | A900 | Received if an invalid query retrieve level is sent. Message shown to user And remains in communication log. |
|  | Unable to process | C000 | Association closed. Message shown to user |

| Service Status | Further Meaning | Status Codes | Reason |
| --- | --- | --- | --- |
|  |  |  | And remains in communication log. |
| Cancel | Matching terminated due to Cancel | FE00 | Association closed. Message shown to user And remains in communication log |
| Pending | Matches are continuing - Current Match is supplied and any Optional Keys were supported in the same manner as Required keys | FF00 | Received. Query is continued |
|  | Matches are continuing - warning that one or more Optional Keys were not supported for existence for this identifier | FF01 | Received. Query is continued |
| Success | Matching is complete - No final Identifier is supplied | 0000 | Association closed |

# MOVE-SCP

#### SOP Classes

MOVE-SCP provides Standard Conformance to the following SOP Classes:

#### Table 4.2.9.1

SOP CLASSES SUPPORTED BY MOVE-SCP AND MOVE-SCU

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Study Root Query/Retrieve Information Model – MOVE | 1.2.840.10008.5.1.4.1.2.2.2 | Yes | Yes |
| Patient Root Query/Retrieve Information Model – MOVE | 1.2.840.10008.5.1.4.1.2.1.2 | No | Yes |
| Patient/Study Only Query/Retrieve Information Model - MOVE | 1.2.840.10008.5.1.4.1.2.3.2 | No | Yes |

### Transfer Syntax Support

For DIMSE C-MOVE operations, Omega AI supports the following transfer syntaxes for C-MOVE command and response messages:

#### Table 4.2.9.2-1

#### Supported Transfer Syntaxes for MOVE-SCP (Command/Response)

| #### Transfer Syntax | #### UID | #### Supported |
| --- | --- | --- |
| #### Implicit VR Little Endian | #### 1.2.840.10008.1.2 | #### Yes |
| #### Explicit VR Little Endian | #### 1.2.840.10008.1.2.1 | #### Yes |
| #### Deflated Explicit VR Little Endian | #### 1.2.840.10008.1.2.1.99 | #### Not proposed |
| #### Explicit VR Big Endian | #### 1.2.840.10008.1.2.2 | #### Not proposed |

Rationale: See Section 4.2.7.2 for rationale on transfer syntax restrictions applicable to all DIMSE Query/Retrieve operations.

Note — C-MOVE Command vs. Image Transfer: The transfer syntax restrictions above apply only to the C-MOVE command and response messages. The transfer syntaxes used for the actual image transfer (C-STORE sub-operations) are negotiated independently based on the supported Storage SOP Classes of the destination AE and are not affected by this limitation.

#### Association Policies

        - General

MOVE-SCP accepts but never initiates associations.

#### Table 4.2.9.3.1

MAXIMUM PDU SIZE FOR MOVE-SCP

| Maximum PDU size received | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

Each time the MOVE-SCP accepts an association, an available thread is assigned to complete the find (query) request. The maximum number of concurrent associations supported by PowerServer is set by configuration. Once this number is reached, no more associations will be accepted until one of the threads become available.

#### Table 4.2.9.3.2

NUMBER OF ASSOCIATIONS AS A SCP FOR MOVE-SCP

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

        - Asynchronous Nature

MOVE-SCP will only allow a single outstanding operation on an Association.

#### Association Initiation Policy

MOVE-SCP initiates an association with the AE specified as the Move Destination in the MOVE request, in order to store the requested instances. The remote AE Title must be configured in Station List.

#### Association Acceptance Policy

When MOVE-SCP accepts an association, it will respond to move requests. The association will be rejected if the Calling AE Title is not configured in the Station List.

        - Activity – Receive Move Request

          - Description and Sequencing of Activities

As requests are received, a STORAGE-SCU operation is initiated to send the requested instances to the specified remote AE.

          - Acceptable Presentation Contexts

#### PRESENTATION CONTEXTS FOR MOVE-SCP AND MOVE-SCU Table 4.2.9.5.1.2

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.9.1 | See Table 4.2.9.1 | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP, SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP, SCU | None |
|  |  | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP, SCU | None |
|  |  | Explicit VR Big Endian | 1.2.840.10008.1.2.2 | SCP, SCU | None |
|  |  | JPEG Lossy Baseline 8-bit Image Compression | 1.2.840.10008.1.2.4.50 | SCP, SCU | None |
|  |  | JPEG Lossy Extended 12-bit Image Compression | 1.2.840.10008.1.2.4.51 | SCP, SCU | None |
|  |  | JPEG Lossless Non-hierarchical (Process 14) | 1.2.840.10008.1.2.4.57 | SCP, SCU | None |

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
|  |  | JPEG Lossless, Non-hierarchical First-Order Prediction (Process 14) | 1.2.840.10008.1.2.4.70 | SCP, SCU | None |
|  |  | JPEG 2000 Lossless Only | 1.2.840.10008.1.2.4.90 | SCP, SCU | None |
|  |  | JPEG 2000 | 1.2.840.10008.1.2.4.91 | SCP, SCU | None |
|  |  | RLE Lossless Transfer Syntax | 1.2.840.10008.1.2.5 | SCP, SCU | None |

            - Extended Negotiation

No extended negotiation is performed. In particular, relational retrievals are not supported.

          - SOP Specific Conformance

            - SOP Specific Conformance to C-MOVE SOP Classes

MOVE-SCP provides standard conformance to the supported C-MOVE SOP Classes. The Patient Root, Study Root, Patient / Study Root only information models are supported. The move operation is performed to the destination AE Title specified in the C-Move request.

Table 4.2.9.5.1.3.1-1

### PATIENT ROOT AND PATIENT / STUDY ONLY REQUESTS IDENTIFIER FOR MOVE-SCP

| Name | Tag | Unique, Matching or Return Key |
| --- | --- | --- |
| PATIENT Level |  |  |
| Patient’s ID | (0010,0020) | U |
| STUDY Level |  |  |
| Study Instance UID | (0020,000D) | U |
| SERIES Level |  |  |
| Study Instance UID | (0020,000D) | U |
| Series Instance UID | (0020,000E) | U |
| IMAGE Level |  |  |
| Study Instance UID | (0020,000D) | U |
| Series Instance UID | (0020,000E) | U |
| SOP Instance UID | (0008,0018) | U |

#### Table 4.2.9.5.1.3.1-2

STUDY ROOT REQUEST IDENTIFIER FOR MOVE-SCP

| Name | Tag | Unique, Matching or Return Key |
| --- | --- | --- |
| STUDY Level |  |  |
| Study Instance UID | (0020,000D) | U |
| SERIES Level |  |  |
| Study Instance UID | (0020,000D) | U |
| Series Instance UID | (0020,000E) | U |
| IMAGE Level |  |  |
| Study Instance UID | (0020,000D) | U |
| Series Instance UID | (0020,000E) | U |
| SOP Instance UID | (0008,0018) | U |

            - Presentation Context Acceptance Criterion

MOVE-SCP will accept the Presentation Contexts given in Table 4.2.9.5.1.2

            - Transfer Syntax Selection Policies

PowerServer transfer syntax selection policy can be configured in the Station List. The support for any of Uncompressed, Lossy Compressed and Lossless Compressed transfer syntaxes may be enabled or disabled. The preference of each type of compression may also be set in the Station List.

            - Response Status

Table 4.2.9.5.1.3.4

### RESPONSE STATUS FOR MOVE-SCP AND SEND TO REMOTE QUERY REQUEST

| Service Status | Further Meaning | Status Codes | Behavior |
| --- | --- | --- | --- |
| Failure | Refused: Out of Resources – Unable to calculate Number of matches | A701 | Unable to calculate number of matches. Association aborted |
| Failure | Refused: Out of resources | A702 | The selected retrieve objects are not online, and cannot be retrieved at this time. |

| Service Status | Further Meaning | Status Codes | Behavior |
| --- | --- | --- | --- |
|  | Unable to perform sub-operations |  | Association aborted |
|  | Refused: SOP Class requested in Association is not supported. | A800 | Requested SOP Class is not Supported Association aborted. |
|  | Refused: Move Destination is unknown Association is not supported. | A801 | Move Destination is unknown. Association aborted. |
|  | Identifier does not match SOP Class | A900 | Sent if an invalid query retrieve level is set |
|  | Unable to process | C000 | Sent if internal database query is Unsuccessful. |
| Cancel | Matching terminated due to Cancel | FE00 | Sent when a C-CANCEL Request is received |
| Warning | Sub-operations Complete – One or more Failures or Warnings | B000 | Sent to initiated AE Title |
| Pending | Sub-operations are continuing | FF00 | Sent for every C-STORE operation completed by STORAGE-SCU. |
| Success | Sub-operations Complete - No Failures | 0000 | Sent to initiated AE Title |

            - Sub-operation behavior

The C-MOVE operation is dependent on completion of C-STORE sub-operations that are occurring on a separate association, which defines transfer syntax selection policies for C-STORE operation. MOVE-SCP initiates a C-STORE sub-operation for each request. The responses from the MOVE-SCP are dependent on the success or failure of the C-STORE sub-operations, not on any explicit action by MOVE-SCP. The remote AE accepts the C-STORE sub-operations, that beyond of the control of MOVE-SCU. If the association on which the C-MOVE was issued is aborted for any reason, the C-STORE sub-operations will continue. If the C-MOVE operation is canceled by the remote AE, MOVE-SCP will also attempt to cancel the corresponding C-STORE sub-operation.

# MOVE-SCU

#### SOP Classes

MOVE-SCU provides Standard Conformance to the SOP Classes shown in Table 4.2.9.1

### Transfer Syntax Support

#### For DIMSE C-MOVE operations, Omega AI supports the following transfer syntaxes for C-MOVE command and response messages:

Table 4.2.10.2-1

Supported Transfer Syntaxes for MOVE-SCU (Command/Response)

| Transfer Syntax | UID | Supported |
| --- | --- | --- |
| Implicit VR Little Endian | 1.2.840.10008.1.2 | Yes |
| Explicit VR Little Endian | 1.2.840.10008.1.2.1 | Yes |
| Deflated Explicit VR Little Endian | 1.2.840.10008.1.2.1.99 | Not proposed |
| Explicit VR Big Endian | 1.2.840.10008.1.2.2 | Not proposed |

Rationale: See Section 4.2.7.2 for rationale on transfer syntax restrictions applicable to all DIMSE Query/Retrieve operations.

Note — C-MOVE Command vs. Image Transfer: The transfer syntax restrictions above apply only to the C-MOVE command and response messages. The transfer syntaxes used for the actual image transfer (C-STORE sub-operations) are negotiated independently based on the supported Storage SOP Classes of the destination AE and are not affected by this limitation.

#### Association Policies

        - General

MOVE-SCU initiates but never accepts associations.

#### Table 4.2.10.3.1

MAXIMUM PDU SIZE RECEIVED AS A SCP FOR MOVE-SCU

| Maximum PDU size received | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

Table 4.2.10.3.2

### NUMBER OF ASSOCIATIONS AS A SCP FOR FIND-SCU

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

        - Asynchronous Nature

MOVE-SCU will only allow a single outstanding operation on an Association.

#### Association Acceptance Policy

MOVE-SCU does not accept associations.

#### Association Initiation Policy

MOVE-SCU attempts to initiate a new association when the user performs the retrieve action from the user interface. (Advanced Search > Remote Studies > Request).

        - Activity – Retrieve Instances from Remote AE

          - Description and Sequencing of Activities

The attempts will be made to retrieve instances from remote AE.

          - Proposed Presentation Contexts

Proposed Presentation Contexts for MOVE-SCU are identical to Accepted Presentation Contexts of MOVE-SCP and illustrated in Table 4.2.9.5.1.2

            - Extended Negotiation

No extended negotiation is performed. In particular, relational retrievals are not supported.

          - SOP Specific Conformance

            - SOP Specific Conformance to C-MOVE SOP Class

#### MOVE-SCU provides standard conformance to the supported C-FIND SOP Classes. Only Study Root Query Information Model is supported and is used at the STUDY level. The requested attributes are illustrated in Table 4.2.10.4.1.3.1

Table 4.2.10.5.1.3.1

STUDY ROOT REQUEST IDENTIFIER FOR FIND-SCU

| Name | Tag | Types of Matching |
| --- | --- | --- |
| STUDY Level |  |  |
| Study Instance UID | (0020,000D) | UNIQUE |

            - Presentation Context Acceptance Criterion

MOVE-SCU will accept the Presentation Contexts given in Table 4.2.9.5.1.2.

            - Transfer Syntax Selection Policies

MOVE-SCU will prefer explicit transfer syntaxes over implicit transfer syntaxes.

## GET-SCP and GET-SCU

#### SOP Classes and Presentation Contexts

GET-SCP and GET-SCU provide Standard Conformance to the following SOP Classes:

Table 4.2.11.1

### SOP CLASSES SUPPORTED BY GET-SCP AND GET-SCU

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Study Root Query/Retrieve Information Model – GET | 1.2.840.10008.5.1.4.1.2.2.2 | Yes | Yes |
| Patient Root Query/Retrieve Information Model – GET | 1.2.840.10008.5.1.4.1.2.1.2 | No | Yes |
| Patient/Study Only Query/Retrieve Information Model - GET | 1.2.840.10008.5.1.4.1.2.3.2 | No | Yes |

#### Table 4.2.11.2

Presentation Contexts for GET-SCP AND GET-SCU

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.11.1 | See Table 4.2.11.1 | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP, SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP, SCU | None |
|  |  | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP, SCU | None |
|  |  | Explicit VR Big Endian | 1.2.840.10008.1.2.2 | SCP, SCU | None |
|  |  | JPEG Lossy Baseline 8-bit Image Compression | 1.2.840.10008.1.2.4.50 | SCP, SCU | None |

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
|  |  | JPEG Lossy Extended 12-bit Image Compression | 1.2.840.10008.1.2.4.51 | SCP, SCU | None |
|  |  | JPEG Lossless Non-hierarchical (Process 14) | 1.2.840.10008.1.2.4.57 | SCP, SCU | None |
|  |  | Explicit VR, Lossless JPEG, Non-hierarchical First-Order Prediction (Process 14) | 1.2.840.10008.1.2.4.70 | SCP, SCU | None |
|  |  | JPEG 2000 Lossless | 1.2.840.10008.1.2.4.90 | SCP, SCU | None |
|  |  | JPEG 2000 | 1.2.840.10008.1.2.4.91 | SCP, SCU | None |
|  |  | RLE Lossless Transfer Syntax | 1.2.840.10008.1.2.5 | SCP, SCU | None |

### Transfer Syntax Support

#### For DIMSE C-GET operations, Omega AI supports the following transfer syntaxes for command and identifier datasets:

#### Table 4.2.11.3-1

Supported Transfer Syntaxes for GET-SCP and GET-SCU (Command/Response)

| Transfer Syntax | UID | Supported |
| --- | --- | --- |
| Implicit VR Little Endian | 1.2.840.10008.1.2 | Yes |
| Explicit VR Little Endian | 1.2.840.10008.1.2.1 | Yes |
| Deflated Explicit VR Little Endian | 1.2.840.10008.1.2.1.99 | Not proposed |
| Explicit VR Big Endian | 1.2.840.10008.1.2.2 | Not proposed |

Rationale: See Section 4.2.7.2 for rationale on transfer syntax restrictions applicable to all DIMSE Query/Retrieve operations.

#### Association Policies

Association Policies for GET-SCP and GET-SCU are identical to MOVE-SCP and MOVE-SCU.

The functional behavior of GET-SCP is identical to MOVE-SCP. The functional behavior of GET-SCU is identical to MOVE-SCU. The differences between those classes that GET-SCU and GET-SCP perform retrieve operations in same association as retrieve requests.

#### SOP Specific Conformance

        - SOP Specific Conformance to C-GET SOP Classes

GET-SCP and GET-SCU provide standard conformance to the supported C-GET SOP Classes. The move operation is performed to the destination AE Title that initiated C-GET request (GET-SCU AE Title).

# MWL-SCP

#### SOP Classes and Presentation Contexts

MWL-SCP provides Standard Conformance to the following SOP Class:

#### Table 4.2.12.1

SOP CLASSES SUPPORTED BY MWL-SCP AND MWL-SCP AND MWL-SCU

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| 1.2.840.10008.5.1.4.31 | 1.2.840.10008.5.1.4.31 | Yes | Yes |

#### Association Policies

        - General

MWL-SCP accepts but never initiates associations.

        - Number of Associations

Each time the MWL-SCP accepts an association, an available thread is assigned to complete the find (query) request. The maximum number of concurrent associations supported by RamSoft PACS is set by configuration. Once this number is reached, no more associations will be accepted until one of the threads become available.

Table 4.2.12.2.2

### NUMBER OF ASSOCIATIONS AS A SCP FOR MWL-SCP

Configurable Parameter

Maximum number of simultaneous associations

#### Association Initiation Policy

MWL-SCP does not initiate associations.

#### Association Acceptance Policy

When MWL-SCP accepts an association, it will respond to query requests. If the Called AE Title does not match the pre-configured AE Title in Station List and option “Allow receiving without verifying AE Title credentials” (See Paragraph 4.4.2Error! Reference source not found.) is unchecked the association will be rejected.

        - Activity – Receive Query Request

          - Description and Sequencing of Activities

As requests are received, the Database is queried for result set.

          - Accepted Presentation Contexts

#### Table 4.2.12.4.1.2

PRESENTATION CONTEXTS FOR MWL-SCP AND MWL-SCU

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.12.1 | See Table 4.2.12.1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP, SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP, SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP, SCU | None |

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance

MWL-SCP provides standard conformance to the supported C-FIND SOP Classes. The maximum number of records to return is identical to number of records for FIND-SCP (See Section 4.2.7.5.1.3.1).

              - The types of Matching supported by the FIND-SCP and FIND-SCU

The types of Matching supported by MWL-SCP and MWL-SCU are identical to types of matching for FIND-SCP and FIND-SCU. (See Section 4.2.7.5.1.3.1.1).

              - Option for return requested attributes

This option is common for FIND-SCP and MWL-SCP (See Section 4.2.7.5.1.3.1.2).

              - MWL Query Information Model.

When the checkbox “Return C-FIND Request Attributes Only” is selected, all requested attributes are returned. In

addition, the following set of attributes is always returned as shown in Table 4.2.12.4.1.3.1.3-1

Table 4.2.12.4.1.3.1.3-1

### SET OF RETURNED ATTRIBUTES FOR MWL-SCP

#### With selected option “Return C-FIND Request Attributes Only”

| Name | Tag | Returned Attribute |
| --- | --- | --- |
| Retrieve AETitle | (0010,0054) | YES |
| Patient’s Name | (0010,0010) | YES |
| Patient’s ID | (0010,0020) | YES |
| Requested Procedure Code Sequence | (0032,1064) | YES |
| Scheduled Procedure Step Sequence | (0040,0100) | YES |
| > Modality | (0008,0060) | YES |
| > Scheduled Station AETitle | (0040,0001) | YES |
| > Scheduled Procedure Step Start Date | (0040,0002) | YES |
| > Scheduled Procedure Step Start Time | (0040,0003) | YES |
| > Scheduled Performing Physician Name | (0040,0006) | YES |
| > Scheduled Protocol Code Sequence | (0040,0008) | YES |

When the checkbox “Return C-FIND Request Attributes Only” is unselected, all requested attributes and additional

attributes are returned as shown in Table 4.2.12.4.1.3.1.3-2

Table 4.2.12.4.1.3.1.3-2

### SET OF RETURNED ATTRIBUTES FOR MWL-SCP

#### With option “Return C-FIND Request Attributes Only” unselected

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Accession Number | (0008,0050) | S,*,U | YES |
| Retrieve AETitle | (0008,0054 | S,*,U | YES |
| Institution’s Name | (0008,0080) | S,*,U | YES |
| Retrieve AETitle | (0008,0054) | NONE | YES |
| Referring Physician’s Name | (0008,0090) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| Referenced Study Sequence | (0008,1110) | NONE | YES |
| >Referenced SOPClassUID | (0008,1150) | NONE | YES |
| >Referenced SOPInstanceUID | (0008,1155) | NONE | YES |
| Referenced Patient Sequence | (0008,1120) | NONE | YES |
| >Referenced SOPClassUID | (0008,1150) | NONE | YES |
| >Referenced SOPInstanceUID | (0008,1150) | NONE | YES |
| Patient’s Name | (0010,0010) | S,*,U | YES |
| Patient’s ID | (0010,0020) | S,*,U | YES |
| Patient’s Issuer | (0010,0021) | S,*,U | YES |
| Patient's Birthdate | (0010,0030) | NONE | YES |
| Patient's Sex | (0010,0040) | NONE | YES |
| Other Patient's ID's | (0010,1000) | NONE | YES |
| Patient’s Size | (0010,1020) | NONE | YES |
| Patient’s Weight | (0010,1030) | NONE | YES |
| Medical Alerts | (0010,2000) | NONE | YES |
| Allergies | (0010,2110) | NONE | YES |
| Additional Patient’s History | (0010,21b0) | NONE | YES |
| Pregnancy Status | (0010,21c0) | NONE | YES |
| Last Menstrual Date | (0010,21d0) | NONE | YES |
| Patient’s Comments | (0010,4000) | NONE | YES |
| Study ID | (0020,0010) | S,*,U | YES |
| StudyInstanceUID | (0020,000D) | UNIQUE | YES |
| RETIRED StudyStatusID | (0032,000A) | S,U | YES |
| RETIRED StudyPriorityID | (0032,000C) | S,U | YES |
| Requesting Physician | (0032,1032) | NONE | YES |
| Requested Procedure Description | (0032,1060) | NONE | YES |
| Requested Procedure Code Sequence | (0032,1064) | NONE | YES |
| >Code Value | (0008,0100) | NONE | YES |
| >Coding Scheme Designator | (0008,0102) | NONE | YES |

| Name | Tag | Types of Matching | Returned Attribute |
| --- | --- | --- | --- |
| >Code Meaning | (0008,0104) | NONE | YES |
| AdmissionID | (0038,0010) | NONE | YES |
| Special Needs | (0038,0050) | NONE | YES |
| Current Patient’s Location | (0038,0300) | NONE | YES |
| Patient’s State | (0038,0500) | NONE | YES |
| Scheduled Procedure Step Sequence | (0040,0100) | NONE | YES |
| >Modality | (0008,0060) | S,U | YES |
| >Scheduled Station AETitle | (0040,0001) | U | YES |
| >Scheduled Procedure Step Start Date | (0040,0002) | S,U,R | YES |
| >Scheduled Procedure Step Start Date | (0040,0003) | NONE | YES |
| >Scheduled Performing Physician’s Name | (0040,0007) | NONE | YES |
| >Scheduled Protocol Code Sequence | (0040,0008) | NONE | YES |
| >>Code Value | (0008,0100) | NONE | YES |
| >>Coding Scheme Designator | (0008,0102) | NONE | YES |
| >>Code Meaning | (0008,0104) | NONE | YES |
| >Scheduled Procedure StepID | (0040,0009) | NONE | YES |
| >Scheduled Station Name | (0040,0010) | NONE | YES |
| >Scheduled Procedure Step Location | (0040,0011) | NONE | YES |
| >Comments On The Scheduled Procedure Step | (0040,0400) | NONE | YES |
| Requested ProcedureID | (0040,1001) | NONE | YES |
| Requested Procedure Priority | (0040,1003) | NONE | YES |
| Patient Transport Arrangements | (0040,1004) | NONE | YES |
| Confidentiality Constraint On PatientData | (0040,3001) | NONE | YES |

            - Presentation Context Acceptance Criterion

MWL-SCP will accept the Presentation Contexts given in Table 4.2.1.4.1.2

            - Transfer Syntax Selection Policies

MWL-SCP will prefer explicit transfer syntaxes over implicit transfer syntaxes. Duplicate Presentation Contexts will not be accepted. If it is offered three identical Presentation Contexts, each of which offers any of the three acceptable

Transfer Syntaxes, it will accept all Presentation Contexts, but with different Transfer Syntaxes in each. Other transfer syntax policies may be configured in the Station List.

            - Response Status

Response Status for MWL-SCP is identical to response status of FIND-SCP as shown in Table HYPERLINK \l "_bookmark71"4.2.7.5.1.3.4

# MWL-SCU

#### SOP Classes

MWL-SCU provides Standard Conformance to the SOP Classes shown in Table 4.2.12.1

#### Association Policies

        - General

FIND-SCU initiates but never accepts associations.

#### Table 4.2.13.2.1

MAXIMUM PDU SIZE FOR MWL-SCU

Configurable Parameter limited to 128 KB

Maximum PDU size

        - Number of Associations

#### Table 4.2.13.2.2

NUMBER OF ASSOCIATIONS FOR MWL-SCU

Configurable Parameter

Maximum number of simultaneous associations

        - Asynchronous Nature

MWL-SCU will only allow a single outstanding operation on an Association.

#### Association Acceptance Policy

MWL-SCU does not accept associations.

#### Association Initiation Policy

MWL-SCU attempts to initiate a new association when the user performs the query action from the user interface. (Advanced Search -> Remote Worklist).

        - Activity – MWL Query Remote AE

          - Description and Sequencing of Activities

The attempt will be made to query the remote AE.

          - Proposed Presentation Contexts

Proposed Presentation Contexts for MWL-SCU are identical to Accepted Presentation Contexts of MWL-SCP and illustrated in Table 4.2.12.4.1.2

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance

            - SOP Specific Conformance to MWL C-FIND SOP Class

MWL-SCU provides standard conformance to the supported C-FIND SOP Classes. The requested attributes are illustrated in Table 4.2.13.4.1.3.1

#### MWL REQUEST Identifier for MWL-SCU Table 4.2.13.4.1.3.1

| Name | Tag | Types of Matching |
| --- | --- | --- |
| Patient’s ID | (0010,0020) | S,*,U |
| Patient’s Issuer | (0010,0021) | S,*,U |
| Patient’s Name | (0010,0010) | S,*,U |
| Modalities In Study | (0008,0061) | S,*,U |
| Institution Name | (0008,0080) | U,* |
| Study Date | (0008,0020) | U,R |
| Study ID | (0020,0010) | S,*,U |
| Accession Number | (0008,0050) | U,* |
| RETIRED StudyStatusID | (0032,000A) | S,*,U |
| RETIRED StudyPriorityID | (0032,000C) | S,*,U |
| RETIRED Study Comments | (0032,4000) | NONE |
| Scheduled AETitle | (0040,0001) | U |

            - Presentation Context Acceptance Criterion

MWL-SCU will accept the Presentation Contexts given in Table 4.2.12.4.1.2.

            - Transfer Syntax Selection Policies

MWL-SCU will prefer explicit transfer syntaxes over implicit transfer syntaxes.

            - Response Status

MWL-SCU will behave as described in Table 4.2.8.5.1.3.4 in response to the status returned in the MWL C-FIND response command messages.

## STORAGE-SCP-MPPS and STORAGE-SCU-MPPS

#### SOP Classes and Presentation Contexts

Table 1.1.1.1-1

### SOP CLASSES SUPPORTED BY STORAGE-SCP-MPPS AND STORAGE-SCU-MPPS

| SOP Class Name | SOP Class UID | SCU | SCP |
| --- | --- | --- | --- |
| Modality Performed Procedure Step SOP Class | 1.2.840.10008.3.1.2.3.3 | Yes | Yes |

#### Table 1.1.1.1-2

PRESENTATION CONTEXTS FOR STORAGE-SCP-MPPS AND STORAGE-SCU-MPPS

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 1.1.1.1-1 | See Table 1.1.1.1-1 | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCP,SCU | None |
|  |  | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCP,SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCP,SCU | None |

#### SOP Specific Conformance

        - SOP Specific Conformance to STORAGE-MPPS Classes

STORAGE-SCP-MPPS and STORAGE-SCU-MPPS provide standard conformance to the supported MPPS SOP Classes. When the association from STORAGE-SCU is initiated, STORAGE-SCU-MPPS performs N-CREATE operation and sets PPSSTATUS of MPPSLIST table to status “STARTED”. After transmission is finished, N-SET operation is invoked and updates above mentioned status to “COMPLETED”.

# PRINT-SCU

#### SOP Classes and Presentation Contexts

PRINT-SCU provides Standard Conformance to the following SOP Classes:

Table 4.2.15.1

### SOP CLASSES SUPPORTED BY PRINT-SCU

| SOP Class Name | SOP Class UID |
| --- | --- |
| Basic Grayscale Print Management Meta SOP Class | 1.2.840.10008.5.1.1.9 |
| Basic Color Print Management Meta SOP Class | 1.2.840.10008.5.1.1.18 |

#### Association Policies

        - General

PRINT-SCU initiates but never accepts associations.

Table 4.2.15.2.1

### MAXIMUM PDU SIZE SENDED FOR PRINT-SCU

| Maximum PDU size sanded | Configurable Parameter limited to 128 KB |
| --- | --- |

        - Number of Associations

Table 4.2.15.2.2

### NUMBER OF ASSOCIATIONS FOR PRINT-SCU

| Maximum number of simultaneous associations | Configurable Parameter |
| --- | --- |

        - Asynchronous Nature

PRINT-SCU will only allow a single outstanding operation on an Association

#### Association Initiation Policy

PRINT-SCU attempts to initiate a new association when User sending images to remote DICOM printer.

        - Activity – Send Print Request

          - Description and Sequencing of Activities

An association is initiated with the selected DICOM printer. The user can select the desired images and the output format. All image display parameters and annotations are burned into the image and the preformatted images are sent to the printer in the background. Any error messages from the printer are recorded in the Print List.

Images are printed as 8 bits or 16 bits per pixel in grayscale or color format.

          - Proposed Presentation Contexts

Table 4.2.15.3.1.2

### PROPOSED PRESENTATION CONTEXTS FOR PRINT-SCU

| Presentation Context Table |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Abstract Syntax |  | Transfer Syntax |  | Role | Extended Negotiation |
| Name | UID | Name | UID |  |  |
| See Table 4.2.15.1 | See Table 4.2.15.1 | Explicit VR Little Endian | 1.2.840.10008.1.2.1 | SCU | None |
|  |  | Implicit VR Little Endian | 1.2.840.10008.1.2 | SCU | None |
|  |  | Deflated Explicit VR, Little Endian | 1.2.840.10008.1.2.1.99 | SCU | None |

            - Extended Negotiation

No extended negotiation is performed.

          - SOP Specific Conformance – Basic Grayscale and Color Print Management

This AE provides standard conformance to this SOP Class as an SCU. All mandatory elements are set for film sessions, film boxes and image boxes. The mandatory attribute, Image Display Format (2010, 0010), has possible values: “STANDARD\x,y” where x,y (columns, rows) combinations can be configured for each target printer. The optional attributes listed in Table 4.2.15.3.1.1-1 can also be set when creating film sessions.

#### Table 4.2.15.3.1.3 -1

Optional Attributes set for Creating Film Sessions

| Entity Name | Tag | Values |
| --- | --- | --- |
| Number Of Copies | (2000,0010) | 1-99 |
| Print Priority | (2000,0020) | LOW, MED, HIGH |
| Medium Type | (2000,0030) | BLUE FILM, CLEAR FILM, PAPER |
| Film Destination | (2000,0040) | MAGAZINE, PROCESSOR |

Table 4.2.15.3.1.3 -2

Optional Attributes set for Creating Film Boxes

| Entity Name | Tag | Values |
| --- | --- | --- |
| Film Orientation | (2010,0040) | Sent only if specified by user PORTRAIT, LANDSCAPE |
| Film Size ID | (2010,0050) | Sent only if specified by user 8INX10IN, 10INX12IN, 10INX14IN, 11INX14IN, 14INX14IN, 14INX17IN, 24CMX24CM, 24CMX30CM, 8_5INX11IN, A4, A3, A |
| Min Density | (2010,0120) | Sent only if specified by user 0 |
| Max Density | (2010,0130) | Sent only if specified by user 399 |

Table 4.2.15.3.1.3-3

Optional Attributes set for Creating Image Boxes

| Entity Name | Tag | Values |
| --- | --- | --- |
| Requested Image Size | (2020,0030) | Values depend on Image Orientation, Film Width/Height/Columns |

## WADO Specifications

WADO Specifications complies with Chapter 9 of PS3.18 Web Services.

#### WADO-URL Specifications

        - WADO – URI Retrieve Rendered Imaging Document

Table 4.2.16.1.1-1

WADO-URL Supported MIME Types

| MIME Type | Description |
| --- | --- |
| application/dicom | The retrieved DICOM object will be encoded correspond to transfer syntax parameter in query. By default, the transfer syntax will be “Explicit VR Little Endian”. |
| image/jpeg | The body content will be a JPEG Image. |

Table 4.2.16.1.1-2

WADO-URL Supported Query Parameters

| Parameter | Kind of Prameter | Description |
| --- | --- | --- |
| requesType | REQUIRED | The value should be “WADO”. |
| studyUID | REQUIRED | The Study Instansce UID of the requested DICOM Object. |
| seriesUID | REQUIRED | The Series Instansce UID of the requested DICOM Object. |
| objectUID | REQUIRED | The SOP Instansce UID of the requested DICOM Object. |
| contentType | OPTIONAL | Possible valuses: “application/dicom” or “ “image/jpeg”. Default is “image/jpeg”. |
| anonymize | OPTIONAL | Possible vaue is “yes”. Parameter shall only be present if contentType is “application/dicom”. The default behavior is not to be anonymize. |
| transferSyntax | OPTIONAL | Parameter shall only be present if contentType is “application/dicom”. The default Transfer Syntax is “Explicit VR Little Endian”. |
| Rows | OPTIONAL | The height of requested image. Parameter shall not be present if contentType is “application/dicom”. |
| columns | OPTIONAL | The width of requested image. Parameter shall not be present if contentType is “application/dicom”. |

| frameNumber | OPTIONAL | Specifies the single frame with that number within a multi-frame image object that shall be returned. The parameter shall not be present if contentType is “application/dicom”. If parameter is omitted, the first frame of a requested multi-frame object will be returned. |
| --- | --- | --- |
| region | OPTIONAL | The parameter allows selection of a rectangular region of an image matrix to be retrieved. The parameter shall not be present if contentType is “application/dicom”. The value shall be expressed as a list of four positive decimal strings, separated by the “,” character. |
| presentationUID | OPTIONAL | SOP Instance UID of the presentation state storage object to be applied to the image. The parameter shall not be present if contentType is “application/dicom”. |
| presentationSeriesUID | REQUIRED and shall only be present if “presentationUID” is present. | Series Instance UID of the series containing the presentation state storage object to be applied to the image. |
| windowWidth | REQUIRED if “windowCenter” Is present. | Control the contrast of the image. This parameter shall not be present if there is “presentationUID” parameter. The parameter shall not be present if contentType is “application/dicom”. |
| windowCenter | REQUIRED if “windowWidth” Is present. | Control the luminosity of the image. This parameter shall not be present if there is “presentationUID” parameter. The parameter shall not be present if contentType is “application/dicom”. |
| imageQuality | OPTIONAL | The parameter shall not be present if contentType is “application/dicom”, except if the the transferSyntax parameter is present and corresponds to a lossy compression . If the requested MIME type is for image/jpeg this parameter indicates the required quality of the image to be returned within the range 1 – 100. 100 being the best quality. |

Table 4.2.16.1.1-3

WADO-URI Retrieve Rendered Imaging Document

| Parameter | Restrictions |
| --- | --- |
| Transfer Syntaxes Supported | Restricted to transfer syntaxes annotated in Table 4.2.3.4.1.2-1 |
| SOP Class Restrictions | Restricted to SOP classes annotated in Table 4.2.3.1 |
| Size Restriction | 1 GB |
| Rendered formats available | Supports JPEG for IMAGE IODS and non-IMAGE IODS |
| Rows restrictions | 16 – 32767 |
| Columns restrictions | 16 - 32767 |
| Window Width restrictions | Whole Window in the range of image pixel values |
| Window Center restrictions | Greater than 4 and whole window in the range of image pixel values. |
| Region restrictions | None |
| Image Quality restrictions | None |
| Annotation restrictions | None |
| Anonymization | None Supported |
| Compression Available | JPEG Lossy |
| Other restrictions | None |

WADO – URL Connection Policies

          - General

All URL connections are limited to HTTP GET Requests.

          - Number of Conections

Maximum Number of simultaneous HTTP Requests is not limited.

## DICOM WEB Support

DICOM Web support provides client URL access for the following operations:

- Echo (Provides the functionality similar to DICOM ECHO SCU)

- Store (Provides the functionality similar to DICOM STORAGE SCU)

- Find (Provides the functionality similar to DICOM FIND SCU)

- Retrieve (Provides the functionality similar to DICOM RETRIEVE SCU)

#### General Settings for DICOM WEB Support

For all functionalities in DICOM Web support we should set Station Type to “DICOM Web” for Specific Station

In Station List General Service (See Figure 4.4.2-3 – Station List – General Service). There are two options for DICOM Web Station types depending on OAuth flow:

- Client Credentials

- JWT Bearer token

The settings for Client Credentials OAuth flow are shown

In Figure 4.4.2-3-1 (Station List – General Service – Set and Check Connection to DICOM Web using Client Credentials for OAuth Flow). This type of OAuth flow should be used when Client Credentials are known.

The following attributes are specified:

- Server URI

- Authorization URI

- ClientID

- Scopes of the Server

- Client Secret

Those values are specified by the Servers.

The settings for JWT Bearer Token OAuth flow for example Station are shown

In Figure 4.4.2-3-2 (Station List – General Service – Set and Check Connection to DICOM Web using JWT Bearer Token for OAuth Flow). We are using this type of OAuth flow when Server gives API to get Access Token and private key (For Example Google HealthCare Server – https://cloud.google.com/healthcare/docs/dicom)

The following attributes are specified:

- Server URI - https://healthcare.googleapis.com/v1beta1/projects/PROJECT_ID/locations/us-central1/datasets/DATA_SET_NAME/dicomStores/DICOM_STORE_NAME/dicomWeb

In this Server URL:

  - PROEJCT_ID – Project ID

  - locations - us-centra1l - location of Server (could be one of Google’s Locations List)

  - datasets - DATA_SET_NAME – The name of the dataset that you are using

  - dicomStores - DATA_STORE_NAME - The name of the datastore that you are using

- JWT Signing Algorithm - RS256

- Token URI (Audience) - https://oauth2.googleapis.com/token

- Claim Issuer (iss) - CLAIM_ISSUER_NAME@PROJECT_ID.iam.gserviceaccount.com (Service Account email). This is defined at the time of creating Service Account The following roles should be granted to the Service Account:

  - Healthcare DICOM Editor

  - Healthcare DICOM Viewer

- Private key is generated (only) when Service Account is created

- Scopes of the Server - https://www.googleapis.com/auth/cloud-healthcare

#### Echo operation

After the configuration of Station with Station Type “DICOM Web” ( See p.4.2.17.1) the “Test Connection” button in Station List can be used to verify the connection. In this case RESTful connection between Client and Server is established.

The HTTP request includes the following attribute in header “Content-Type”:

- The MIME type “application/dicom”

“DICOM Connection succeeded” message is appeared in case of successful connection. To detail see Figure 4.4.2-3-1 or Figure 4.4.2-3-2 depending on OAuth Flow.

#### Store operation

Store operation is compatible with Store Transaction in PS3.18 Web Services for Client in Study Level.

The HTTP request for store the data incudes the following attributes in header “Content-Type”:

- The multipart/related media type

- The MIME type “application/dicom”

- The boundary separator

The HTTP response Status Codes identical to Status Codes of Chapter 8.5 PS3.18 Web Services.

        - Activity – Send Store Request

Perform the configuration of Station with Station Type “DICOM Web” ( See p.4.2.17.1). Search the Study for send in GUI Search and click button “Send Study to another station”. Choose the “DICOM Web” station and click “OK” button.

The sequence of activity to send the study of the patient “AXL489273” to DICOM Web” station “NHUNG-DICOMWEB” is shown in Figure 4.2.17.3.1-1

Figure 4.2.17.3.1-1  Send the study of the patient “AXL489273” to DICOM Web” station “NHUNG-DICOMWEB”

The following RESTful requests were sent to the servers of the stations

- Station – “NHUNG-DICOMWEB”

https://healthcare.googleapis.com/v1beta1/projects/rapidresults-248519/locations/us-central1/datasets/RSQA/dicomStores/image_volume_1/dicomWeb/studies?StudyInstanceUID=NULL&li mit=1.

In this case StudyInstanceUID is retrieved from binary DICOM file.

- Station – “DicomWeb_MAIN_1B” “https://qa-main-1-dicom-api-

webapp.azurewebsites.net/dicom/studies/1.2.892.999888.11.7.119787.86546.4568777”

#### Query operation

Query operation is compatible with Search Transaction in PS3.18 Web Services for Client in Study Level. The HTTP request for query incudes the following attribute in header “Content-Type”:

- The MIME type “application/dicom+json”

- The multipart/related media type

The attributes of Query Operation are illustrated in Table 4.2.17.4.1

#### Table 4.2.17.4.1

ATTRIBUTES FOR QUERY OPERATION

| Name | Tag | Types of Matching |
| --- | --- | --- |
| Patient’s ID | (0010,0020) | S,*,U |
| Patient’s Issuer | (0010,0021) | S,*,U |
| Patient’s Name | (0010,0010) | S,*,U |
| Patient’s BirthDay | (0010,0030) | S,*,U,R |
| Patient’s Sex | 0010,0040) | S,*,U |
| Other Patient's ID's | (0010,1000) | S,*,U |
| Modalities In Study | (0008,0061) | S,*,U |
| Institution Name | (0008,0080) | U,* |
| Study Date | (0008,0020) | U,R |
| Study Time | (0008,0030) | U,R |
| Study ID | (0020,0010) | S,*,U |
| Accession Number | (0008,0050) | U,* |
| RETIRED StudyStatusID | (0032,000A) | S,*,U |
| RETIRED StudyPriorityID | (0032,000C) | S,*,U |
| StudyInstanceUID | (0020,000D) | S |
| StudyDescription | (0008,1030) | S |

The tag without specified value that should be returned can be added as “includefield”. “fuzzymatching” in query is supported.

“limit” and offset query parameters are supported. The maximum value of “limit” depends of Server.

The HTTP response Status Codes identical to Status Codes of Chapter 8.5 PS3.18 Web Services.

        - Activity – Query Request

Perform the configuration of Station with Station Type “DICOM Web” ( See p.4.2.17.1). Search the Study for query in GUI Advance Search for “DICOM Web” Statiion.

The sequence of activity to query the study of the patient “AXL489273” from DICOM Web” station “NHUNG-DICOMWEB” is shown in Figure 4.2.17.4.1-1

Figure 4.2.17.4.1-1 Query the study of the patient “AXL489273” from DICOM Web” station “NHUNG-DICOMWEB”

- The following RESTful request was sent to the server of the station “NHUNG-DICOMWEB” “https://healthcare.googleapis.com/v1beta1/projects/rapidresults-248519/locations/us-central1/datasets/RSQA/dicomStores/image_volume_1/dicomWeb/studies?fuzzymatching=true&Patie ntName=AXL489273&includefield=00080201&includefield=00080020&includefield=00080030&includef ield=0032000C&includefield=00100021&includefield=00100030&includefield=00100010&includefield= 00100040&includefield=00100020&includefield=00200010&includefield=0020000D&includefield=0008 0050&includefield=00080061&includefield=00081030&includefield=0032000A&includefield=00080080 &includefield=00180015”

All Include fields and their tags were mentioned in 4.2.7.4

#### Retrieve operation

Retrieve operation is compatible with Retrieve Transaction in PS3.18 Web Services for Client in Study Level.

The HTTP request for store the data incudes the following attributes in header “Content-Type”:

- The multipart/related media type

- The MIME type “application/dicom”

- The boundary separator

The HTTP response Status Codes identical to Status Codes of Chapter 8.5 PS3.18 Web Services.

        - Activity – Retrieve Request

Perform the configuration of Station with Station Type “DICOM Web” ( See p.4.2.17.1).

Search the Study for query in GUI Advance Search for “DICOM Web” Statiion.

The sequence of activity to query the study of the patient “AXL489273” from DICOM Web” station “NHUNG-DICOMWEB” is shown in Figure 4.2.17.4.1-1.

After finishing the search from “DICOM Web” station select the study and click “Request selected study from station” button. The selected study is extracted from “DICOM Web” station via WADO request.

## Network Interfaces

## Physical Network Interface

All RamSoft PACS AE Titles provide DICOM V3.0 TCP/IP Network Communication Support as defined in Part 8 of the DICOM Standard.

#### TCP/IP Stack

All RamSoft PACS AE Titles inherit their TCP/IP stack from the Windows system upon which they execute. RamSoft PACS has been tested with Winsock 2.2.2 on Windows 7 32-bit, Windows 7 64-bit, Windows 8 32-bit, Windows 8 64-bit, Windows 8.1 32-bit, Windows 8.1 64-bit, Windows 2008 R2 Server, Windows 2012 Server, Windows 2012 R2 Server.

#### Physical Media Support

All RamSoft PACS AE’s are indifferent to the physical medium over which TCP/IP executes; they inherit this from the Windows system upon which they execute.

## IPv4 and IPv6 support

The Application supports IPv4.

## Configuration

All configuration of the application is performed through the use of configuration files stored in pre-defined locations and in Database.

## AE Title/Presentation Address Mapping

The mapping from AE Title to TCP/IP addresses and ports is configurable through the Station List. Refer to the next Section of this document for more details.

## The main interfaces for configuration settings

The main interface for configuration settings is show in the following figures:

Figure 4.4.2-1  – DICOM Server Configuration

Figure 4.4.2-2 – DICOM Print SCU Configuration

Figure 4.4.2-3 – Station List – General Service

Figure 4.4.2-3-1 – Station List – General Service-Set and Check Connection to DICOM WEB using Client Credentials for OAuth Flow

Figure 4.4.2-3-2 – Station List – General Service-Set and Check Connection to DICOM WEB using JWT Bearer Token for OAuth Flow

Figure 4.4.2-4 – Station List - Send Service

Figure 4.4.2-5 – Station List - Receive Service

Figure 4.4.2-6 – Station List - Query Service

Figure 4.4.2-7 – Station List - Print Service

## Parameters

Table 4.4.3-1 Configuration Parameters Table

| Parameter | Configurable | Default Value |
| --- | --- | --- |
| General Parameters |  |  |
| PDU Size | Yes | 64 KB |
| Timeout for association responses, general DIMSE timeout, and TCP/IP messages | Yes | 60 seconds |
| Any changes to default TCP/IP settings, such as configurable stack parameters. | No | None |
| Server Application AETitle Name | Yes | RAMSOFT |
| Server Station Name | Yes | RAMSOFT |
| Server Application Port | Yes | 104, 2762 (for TLS) |
| Maximum of simultaneous Associations by Service and/or SOP Class | Yes | 10 |
| Encrypt Data (TLS) | Yes | No |
| Server Send Retry Interval | Yes | 10000 (Timer interval in milliseconds at which the server will check for new items in the Transmit, Request, and Print Lists). |
| Capture Format | Yes | 3 (Modifies the SOP Class used for newly created DICOM image objects. 0 - Ultrasound 1 - Secondary Capture(old) 2 - Fluoroscopy 3-Secondary Capture Multi-Frame (new)) |
| Heartbeat Interval | Yes | 600000 (Heartbeat retry interval in milliseconds). |

| Parameter | Configurable | Default Value |
| --- | --- | --- |
| AE Specific Parameters (all AE’s) |  |  |
| Size constraint for maximum object size | No | None |
| Maximum PDU size the AE can send | Yes | 128 KB |
| Maximum PDU size the AE can receive | Yes | 128 KB |
| Maximum concurrent print associations | Yes | 5 |
| Maximum concurrent request associations | Yes | 5 |
| Maximum concurrent transmit associations | Yes | 5 |
| SOP Class support | No | All supported SOP Classes always proposed and accepted |
| Transfer Syntax support | Yes | All supported Transfer Syntaxes |
| YBR Color Photometric Interpretation (when sending) Note: When this option is unselected, color images will be sent in RGB photometric interpretation. | Yes | Yes |
| Burn Presentation State (when sending) | Yes | No |
| Support Storage Commitment (when sending) | Yes | No |
| Support MPPS (when sending) | Yes | No |
| Sent Patient Documents Note: When this option is selected, Patient Level Documents are sent. | Yes | Yes |
| Custom Scripting (when sending) Note: Custom Scripting Button is used to modify DICOM objects before they are sent. | Yes | None |
| Script Execution Timeout | Yes | 5 minutes |
| Return C-FIND Request Attributes Only | Yes | No |
| Set Study Status when Storage Commitment is Received | Yes | Yes |
| Send Event Report Message on the Same Association (For Storage Commitment Response) | Yes | No |
| Ignore Presentation States (when receiving) | Yes | No |
| Receive Study As (Configured Status) | Yes | COMPLETED |
| (Receive Study As Configured) Facility | Yes | No |
| (Receive Study As Configured) Issuer of PatientID | Yes | No |

| Parameter | Configurable | Default Value |
| --- | --- | --- |
| Always Overwrite Status (when receiving) Note: When this option is selected, incoming studies will have the status selected in the Receive Study As field no matter what their previous status was. | Yes | Yes |
| Support MPPS (when receiving) | Yes | No |
| Repair Patient Name (when receiving) Note: When this option is selected, incoming patient names are checked to presence of carets (^). If they are missing, they will automatically be placed into the name. This helps in reducing conflicts and matching up priors with existing patients. | Yes | No |
| Overwrite Existing Objects (when receiving) Note: When this option is selected, existing objects will be replaced with received objects. | Yes | No |
| Custom Scripting (when receiving) Note: Custom Scripting Button is used to modify DICOM receiving objects before they are stored. | Yes | None |
| Use C-GET (when retrieve request sent) Note: When this option is selected, C-GET retrieve requests are used. If DICOM queries are not working properly when attempting to use C-GET, de-selecting this option will force PowerServer to use the C-MOVE request instead. | Yes | Yes |
| Clear Issuer Of Patient ID (when query request sent) Note: When this option is selected, the Issuer of Patient ID will be cleared in C-MOVE requests. | Yes | No |
| Prefer Study Root Query Retrieve Model (when retrieve request sent) Note: When this option is selected, Study Root Query Retrieve Model is used for C-MOVE requests. | Yes | Yes |
| Maximum number of records to return in C-FIND query And MWL query | Yes | 100 |

# MEDIA INTERCHANGE

## Implementation Model

## Application Data Flow

The Media Application Entity exports Images, Presentation States, SR, CADSR, PDF Objects to a CD-R Storage medium. It is associated with the local real-world activity "Write Study to a disk”. " Write Study to a disk " is performed upon user request for selected studies.

## Functional Definition of AEs

#### Functional Definition of Media Application Entity

Activation of the " Write Study to a disk " icon will pass the currently selected studies to the Media Application Entity. The SOP Instances associated with the selection will be collected into export job. The contents of export job will be written to a single medium.

## Sequencing of Real-Word Activities

At least one DICOM object should be selected before the Media Application Entity can be invoked. Table 5.1.3-1 shown the options for selection the objects to write to CD/DVD.

#### Table 5.1.3-1

THE OPTIONS FOR SELECTION OBJECTS TO WRITE TO CD-R

## File Meta Information Options

The implementation information written to the File Meta Header in each file is identical to the information described in Table 4.2.

## AE Specifications

## Media Application Entity Specification

The Media Application Entity provides standard conformance to the Media Storage Service Class. The Application Profiles and roles are listed in Table 1-2

#### File Meta Information for the Application Entity

The Source Application Entity Title included in the File Meta Header is identical to Server Application Entity Title when disc is created (See Paragraph 4.4.3).

#### Real-World Activities

        - Activity – Export to External Media

The Offline-Media Application Entity acts as an FSC when requested to export the studies from the local database to a disc. The dialogue will be presented allowing the user to select the DICOM objects

for export (See par. 5.1.3). The user will be prompted to insert an empty disc for scheduling the export job. The contents of the export job will be written together with a corresponding DICOMDIR to the disc.

The user can cancel an export job in the queue.

        - Activity – Reading from External Media

RamSoft PACS has independent application PRViewer which reading full contents of disc. The description of PRViewer goes out of scope of DICOM Conformance Statement.

In addition the Folder where DICOM objects of disc are located can be imported to RamSoft PACS.

# TRANSFORMATION OF DICOM TO CDA

None supported.

# SUPPORT OF CHARACTER SETS

## Character Sets

In addition to the default character repertoire, the Defined Terms for Specific Character Set in Table 7.1-1 are supported:

#### Table 7.1-1

Supported Specific Character Set Defined Terms

| Character Set Description | Defined Term |
| --- | --- |
| ASCII Default repertoire | ISO_IR 6 (Default) |
| Unicode UTF-8 | ISO_IR 192 |
| Latin alphabet No. 1 | ISO_IR 100 |
| Latin alphabet No. 2 | ISO_IR 101 |
| Latin alphabet No. 3 | ISO_IR 109 |
| Latin alphabet No. 4 | ISO_IR 110 |
| Cyrillic alphabet | ISO_IR 144 |
| Arabic alphabet | ISO_IR 127 |
| Greek alphabet | ISO_IR 126 |
| Hebrew alphabet | ISO_IR 138 |
| Latin alphabet No. 5 | ISO_IR 148 |

| Thai | ISO_IR 166 |
| --- | --- |
| Chinese | GB 18030 |
| Japanese | ISO_IR 13 |

# SECURITY

## Security Profiles

RamSoft PACS supports “Basic TLS Secure Transport Connection Profile” and “AES TLS Secure Transport Connection Profile”. Those profiles are defined in Part 15 of the DICOM Standard for secure DICOM Communication. Table 8.1-1 specifies mechanisms that shall be supported if the corresponding features within TLS are supported by the Application Entity. Transport Layer Security Version 1.0 protocol is used.

#### Table 8.1-1 Supported TLS Feature

| Supported TLS Feature | Mechanism |
| --- | --- |
| Entity Authentication | RSA based certificates |
| Exchange of Master Secrets | RSA |
| Data Integrity | SHA |
| Privacy | CBC, Triple DES EDE |

The list of the ciphersuites for negotiation is shown in Table 8.1-2, with the most preferred ciphersuites on top.

#### Table 8.1-2

List of the ciphersuites for secure negotiation

| The name of ciphersuite | Key Exchange | Encryption | Secure Hash |
| --- | --- | --- | --- |
| TLS_DHE_RSA_WITH_AES_128_GCM_SHA256 | DHE | AES (CBC mode) with 128 bit key | SHA256 |
| TLS_ECDHE_RSA_WITH_AES_128_GCM_SHA256 | ECDHE | AES (CBC mode) with 128 bit key | SHA256 |
| TLS_DHE_RSA_WITH_AES_256_GCM_SHA384 | DHE | AES (CBC mode) with 256 bit key | SHA384 |
| TLS_ECDHE_RSA_WITH_AES_256_GCM_SHA384 | ECDHE | AES (CBC mode) with 256 bit key | SHA384 |
| TLS_RSA_WITH_AES_128_CBC_SHA | RSA | AES (CBC mode) with 128 bit key | SHA1 |
| TLS_RSA_WITH_3DES_EDE_CBC_SHA | RSA | Triple DES (EDE CBC mode) with 168 bit key | SHA1 |

When a TLS integrity check fails, the connection is dropped and an A-P-ABORT indication is issued

to the application by the DICOM upper layer state machine, together with an implementation-specific

provider reason. For the RamSoft PACS implementation the provider reason for any A-P-ABORT indication is

always “unspecified”. The error code reported to the application layer is DUL_PEERABORTEDASSOCIATION.

## Association Level Security

The associations can be established only with DICOM Nodes defined in Station List (To details see the Fig. 4.4.2 - 3 Station List – General Service). A private key and certificate must be generated and peer certificates must be installed for all connected SCUs. Care must be taken while creating and exchanging certificates ensuring the privacy and authenticity of them.

## Application Level Security

Application Level Security is enabled in Station List. The MWL-SCU and FIND-SCU queries from receiving DICOM node can be filtered by Issuer of PatientID, Facility, Study Status. To details see the Fig. 4.4.2 -5 Station List – Receive Service.

# ANNEXES

## IOD Contents

## Created SOP Instances

This section specifies the subsets of DICOM Information Object Definitions (IOD) used to represent the information objects produced by RamSoft PACS.

The abbreviations are used in the below tables are following:

VNAP - Value Not Always Present (attribute sent zero length if no value is present)

ANAP - Attribute Not Always Present ALWAYS - Always Present with a value EMPTY - Attribute is sent without a value

The abbreviations are used for the source of the data values in the tables are:

USER -The attribute value source is from User input

AUTO - The attribute value is generated automatically

MWL, MPPS, STORE SCU - The attribute value is the same as the value received using a DICOM service such as Modality Worklist, Modality Performed Procedure Step, STORE SCU Classes

CONFIG - the attribute value source is a configurable parameter

#### Grayscale Softcopy Presentation State IOD

This section describes the Grayscale Softcopy Presentation State IODs which are created by RamSoft PACS.

Table 9.1.1.1 -1

IOD of Created Grayscale Softcopy Presentation State SOP Instances

| IE | MODULE | REFERENCE | PRESENCE OF MODULE |
| --- | --- | --- | --- |
| Patient | Patient | Table 9.1.1.6.1 - 1 | ALWAYS |
| Visit | Visit Admission Module | Table 9.1.1.6.2 - 1 | ALWAYS |
|  | Visit Identification Module | Table 9.1.1.6.3 - 1 | ALWAYS |
|  | Visit Status Module | Table 9.1.1.6.4 - 1 | ALWAYS |
| Study | General Study | Table 9.1.1.6.5 - 1 | ALWAYS |
|  | Requested Procedure Module | Table 9.1.1.6.6 - 1 | VNAP |
|  | Patient Study | Table 9.1.1.6.7 - 1 | ALWAYS |
| Series | General Series | Table 9.1.1.6.8 - 1 | ALWAYS |
| Equipment | General Equipment | Table 9.1.1.6.9 - 1 | ALWAYS |
| Image | General Image | Table 9.1.1.7.3 - 1 | ALWAYS |
| Presentation State | Presentation State | Table 9.1.1.7.1 - 1 | ALWAYS |
|  | SOP Common | Table 9.1.1.7.2 - 1 | ALWAYS |
|  | Display Shutter | Table 9.1.1.7.4 - 1 | Only if Shutter is present |

|  | Displayed Area | Table 9.1.1.7.5 - 1 | ALWAYS |
| --- | --- | --- | --- |
|  | Graphic Annotation | Table 9.1.1.7.6 - 1 | Only if Graphic Annotations are present |
|  | Graphic Layer | Table 9.1.1.7.7 - 1 | Only if Graphic Annotations are present |
|  | Modality LUT | Table 9.1.1.7.8 - 1 | ALWAYS |
|  | Softcopy VOI LUT | Table 9.1.1.7.9 - 1 | ALWAYS |
|  | Softcopy Presentation LUT | Table 9.1.1.7.10 - 1 | ALWAYS |

#### Secondary Capture IODs

This section describes the Secondary Capture IODs which are created by RamSoft PACS when the user imports the image to Study in Patient Explorer.

Table 9.1.1.2 -1

IOD of Created Secondary Capture SOP Instances

| IE | MODULE | REFERENCE | PRESENCE OF MODULE |
| --- | --- | --- | --- |
| Patient | Patient | Table 9.1.1.6.1 - 1 | ALWAYS |
| Visit | Visit Admission Module | Table 9.1.1.6.2 - 1 | ALWAYS |
|  | Visit Identification Module | Table 9.1.1.6.3 - 1 | ALWAYS |
|  | Visit Status Module | Table 9.1.1.6.4 - 1 | ALWAYS |
| Study | General Study | Table 9.1.1.6.5 - 1 | ALWAYS |
|  | Requested Procedure Module | Table 9.1.1.6.6 - 1 | VNAP |
|  | Patient Study | Table 9.1.1.6.7 - 1 | ALWAYS |
| Series | General Series | Table 9.1.1.6.8 - 1 | ALWAYS |
| Equipment | General Equipment | Table 9.1.1.6.9 - 1 | ALWAYS |
|  | SC Equipment | Table 9.1.1.8.1 - 1 | ALWAYS |
| Image | General Image | Table 9.1.1.8.2 - 1 | ALWAYS |
|  | Image Pixel | Table 9.1.1.8.3 - 1 | ALWAYS |
|  | SOP Common | Table 9.1.1.8.4 - 1 | ALWAYS |
|  | Modality LUT | Table 9.1.1.8.5 -1 | ALWAYS |
|  | Softcopy VOI LUT | Table 9.1.1.8.6 - 1 | ALWAYS |

#### Hardcopy IODs

This section describes the Hardcopy IODs which are created by RamSoft PACS when the user scanning the image to Study in Patient Explorer.

Table 9.1.1.3 -1

IOD of Created Hardcopy SOP Instances

| IE | MODULE | REFERENCE | PRESENCE OF MODULE |
| --- | --- | --- | --- |
| Patient | Patient | Table 9.1.1.6.1 - 1 | ALWAYS |
| Visit | Visit Admission Module | Table 9.1.1.6.2 - 1 | ALWAYS |
|  | Visit Identification Module | Table 9.1.1.6.3 - 1 | ALWAYS |
|  | Visit Status Module | Table 9.1.1.6.4 - 1 | ALWAYS |
| Study | General Study | Table 9.1.1.6.5 - 1 | ALWAYS |
|  | Requested Procedure Module | Table 9.1.1.6.6 - 1 | VNAP |
|  | Patient Study | Table 9.1.1.6.7 - 1 | ALWAYS |
| Series | General Series | Table 9.1.1.6.8 - 1 | ALWAYS |
| Equipment | General Equipment | Table 9.1.1.6.9 - 1 | ALWAYS |
|  | SC Equipment | Table 9.1.1.9.1 - 1 | ALWAYS |
| Image | General Image | Table 9.1.1.9.2 - 1 | ALWAYS |
|  | Image Pixel | Table 9.1.1.9.3 - 1 | ALWAYS |
|  | SOP Common | Table 9.1.1.9.4 - 1 | ALWAYS |
|  | Modality Lut | Table 9.1.1.9.5 - 1 | ALWAYS |
|  | Softcopy VOI LUT | Table 9.1.1.9.6 - 1 | ALWAYS |
|  | Softcopy Presentation LUT Module | Table 9.1.1.9.7 - 1 | ALWAYS |

#### Encapsulated PDF Document IODs

This section describes the PDF Document IODs which are created by RamSoft PACS when the user performs attachment of document in Document Viewer

Table 9.1.1.4 - 1

IOD of PDF Documents SOP Instances

| IE | MODULE | REFERENCE | PRESENCE OF MODULE |
| --- | --- | --- | --- |
| Patient | Patient | Table 9.1.1.6.1 - 1 | ALWAYS |
| Visit | Visit Admission Module | Table 9.1.1.6.2 - 1 | ALWAYS |
|  | Visit Identification Module | Table 9.1.1.6.3 - 1 | ALWAYS |
|  | Visit Status Module | Table 9.1.1.6.4 - 1 | ALWAYS |
| Study | General Study | Table 9.1.1.6.5 - 1 | ALWAYS |
|  | Requested Procedure Module | Table 9.1.1.6.6 - 1 | VNAP |
|  | Patient Study | Table 9.1.1.6.7 - 1 | ALWAYS |
| Series | General Series | Table 9.1.1.6.8 - 1 | ALWAYS |
| Equipment | General Equipment | Table 9.1.1.6.9 - 1 | ALWAYS |
|  | SC Equipment | Table 9.1.1.10.1 - 1 | ALWAYS |
|  | General Image | Table 9.1.1.10.2 - 1 | ALWAYS |

| Encapsulated Document | Encapsulated Document | Table 9.1.1.10.3 - 1 | ALWAYS |
| --- | --- | --- | --- |
|  | SOP Common | Table 9.1.1.10.4 - 1 | ALWAYS |

#### Structured Report Document IODs

This section describes the Structured Report Document IODs which are created by RamSoft PACS when the user creating report or performs attachment of document to the report in Document Viewer. The implementation supports all three general purpose SR document classes as specified in the DICOM 2003 standard: Basic Text SR, Enhanced SR and Comprehensive SR.

Table 9.1.1.5 - 1

IOD of Structured Report Document SOP Instances

| IE | MODULE | REFERENCE | PRESENCE OF MODULE |
| --- | --- | --- | --- |
| Patient | Patient | Table 9.1.1.6.1 - 1 | ALWAYS |
| Visit | Visit Admission Module | Table 9.1.1.6.2 - 1 | ALWAYS |
|  | Visit Identification Module | Table 9.1.1.6.3 - 1 | ALWAYS |
|  | Visit Status Module | Table 9.1.1.6.4 - 1 | ALWAYS |
| Study | General Study | Table 9.1.1.6.5 - 1 | ALWAYS |
|  | Requested Procedure Module | Table 9.1.1.6.6 - 1 | VNAP |
|  | Patient Study | Table 9.1.1.6.7 - 1 | ALWAYS |
| Series | General Series | Table 9.1.1.6.8 - 1 | ALWAYS |
| Equipment | General Equipment | Table 9.1.1.6.9 - 1 | ALWAYS |
| Document | General Image | Table 9.1.1.11.1 - 1 | ALWAYS |
|  | SR Document General | Table 9.1.1.11.2 - 1 | ALWAYS |
|  | SR Document Content (Text Content) | Table 9.1.1.11.3 - 1 | ALWAYS |
|  | SR Document Content (Binary Content) | Table 9.1.1.11.4 - 1 | ALWAYS |
|  | SOP Common | Table 9.1.1.11.5 - 1 | ALWAYS |

#### Common Modules

        - Patient Module

#### Table 9.1.1.6.1- 1

Patient Module of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presenc e of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Patient’s Name | (0010,0010) | PN | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Patient ID | (0010,0020) | LO | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Issuer of Patient ID | (0010,0021) | LO | From Modality WorkList or Store SCU or auto generated | ALWAYS | MWL / STORE SCU / AUTO |
| Patient’s Birth Date | (0010,0030) | DA | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient’s Sex | (0010,0040) | CS | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient's Insurance Plan Code Sequence | (0010,0050) | SQ | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| > Code Value | (0008,0100) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient’s Address | (0010,1040) | CS | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient’s Primary Language Code Sequence | (0010,0101) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Ethnic Group | (0010,2160) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| > Code Value | (0008,0100) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| RamSoft Race Identifier | (3129,0010) | Private Tag (LO) | Copied from referenced Image Object | ALWAYS | AUTO |
| Race | (3129,1010) | Private Tag (LO) | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER / AUTO |
| Other Patient IDs | (0010,1000) | LO | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Country of Residence | (0010,2150) | LO | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient’s Telephone Numbers | (0010,2154) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |

| Smoking Status | (0010,21A0) | CS | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| --- | --- | --- | --- | --- | --- |
| Patient’s Mother’s Birth Name | (0010, 1060) | PN | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Allergies | (0010,2110) | LO | From Modality WorkList or STORE SCU or user Input The format is following: ALLERGY NAME\ALLERGY NOTES\SEVERITY | VNAP | MWL / STORE SCU / USER |
| Patient’s Comments | (0010,4000) | LT | From Store SCU or user Input | VNAP | STORE SCU / USER |

        - Visit Admission Module

#### Table 9.1.1.6.2- 1

Visit Admission Module of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Route of Admissions | (0038,0016) | LO | From Modality WorkList or Store SCU | VNAP | MWL, STORE SCU |
| Admitting Date | (0038,0020) | DA | From Store SCU | VNAP | STORE SCU |
| Admitting Time | (0038,0021) | TM | From Store SCU | VNAP | STORE SCU |

        - Visit Identification Module

#### Table 9.1.1.6.3- 1

Visit Identification Module of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Admission ID | (0038,0010) | LO | From Store SCU | VNAP | STORE SCU |

        - Visit Status Module

#### Table 9.1.1.6.4 - 1

Visit Module Status of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Current Patient Location | (0038,0300) | LO | From Store SCU or user Input | VNAP | STORE SCU / USER |
| Visit Comments | (0038,4000) | LT | From Store SCU or user Input | VNAP | STORE SCU / USER |

        - General Study Module

#### Table 9.1.1.6.5 - 1

General Study Module Attributes of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Study Instance UID | (0020,000D) | UI | From Modality WorkList or Store SCU or auto generated | ALWAYS | MWL / STORE SCU / AUTO |
| Study Date | (0008,0020) | DA | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU |
| Study Time | (0008,0030) | TM | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU |
| Study ID | (0020,0010) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / AUTO |
| Accession Number | (0008,0050) | SH | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Study Description | (0008,1030) | LO | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |

| Referring Physician’s Name | (0008,0090) | PN | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU /USER |
| --- | --- | --- | --- | --- | --- |
| Physician(s) of Record | (0008,1048) Value Multiplicity (1-n) | PN | From Modality WorkList or Store SCU or user Input (List of Consulting Physicians) | VNAP | MWL / STORE SCU / USER |
| Performing Physician’s Name | (0008,1050) | PN | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Name of Physician Reading Study | (0008,1060) | PN | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Interpretation Transcriber | (4008,010a) | PN | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Operator’s Name | (0008,1070) | PN | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Reason For Study | (0032,1030) | LO | From Modality WorkList or Store SCU or user Input (Study Symptoms) | VNAP | MWL / STORE SCU / USER |
| Study Comments | (0032,4000) | LT | From Modality WorkList or Store SCU or user Input (Study Comments) | VNAP | MWL / STORE SCU / USER |
| Study Status ID | (0032,000A) | CS | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Study Priority ID | (0032,000C) | CS | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |
| Procedure Code Sequence | (0008,1032) | SQ | From Modality WorkList or Store SCU or user Input (One or more Items are permitted in this Sequence) | VNAP | MWL / STORE SCU / USER |
| -> Code Value | (0008,0100) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| -> Coding Scheme Designator | (0008,0102) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| -> Code Meaning | (0008,0104) | LO | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |

        - Requested Procedure Module

#### Table 9.1.1.6.6 - 1

Requested Procedure Module Attributes of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Requested Procedure ID | (0040,1001) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Requested Procedure Priority | (0040,1003) | SH | From Modality WorkList or Store SCU or user Input | ALWAYS | MWL / STORE SCU / USER |

        - Patient Study Module

#### Table 9.1.1.6.7 - 1

Patient Study Module Attributes of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Patient’s Height | (0010,1020) | DS | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Patient’s Weight | (0010,1030) | DS | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Additional Patient History | (0010,21B0) | LT | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| Admitting Diagnoses Code Sequence | (0008,1084) | SQ | From Modality WorkList or Store SCU or user Input (One or more Items are permitted in this Sequence) | VNAP | MWL / STORE SCU / USER |
| -> Code Value | (0008,0100) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| -> Coding Scheme Designator | (0008,0102) | SH | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |
| -> Code Meaning | (0008,0104) | LO | From Modality WorkList or Store SCU or user Input | VNAP | MWL / STORE SCU / USER |

        - General Series Module

#### Table 9.1.1.6.8 - 1

General Series Module Attributes of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Modality | (0008,0060) | CS | - PR – for Created Grayscale Softcopy Presentation State SOP Instances - SC – for Created Secondary Capture SOP Instances - HC – for Created Hardcopy SOP Instances - SR – for Created Structured Report SOP Instances - DOC – for Created Encapsulated PDF Documents SOP Instances | ALWAYS | AUTO |
| Series Instance UID | (0020,000E) | UI | Auto Generated Using RamSoft UID prefix “1.2.124.113540.1.3” | ALWAYS | AUTO |
| Series Date | (0008,0021) | DA | Equal to Study Date – for all created SOP Instances | ALWAYS | AUTO |
| Series Time | (0008,0031) | TM | Equal to Study Time – for all created SOP Instances | ALWAYS | AUTO |
| Performing Physician’s Name | (0008,1050) | PN | From Modality WorkList or STORE SCU or user Input One of Performing Physicians of Study | VNAP | AUTO |
| Operator’s Name | (0008,1070) | PN | Technician in Study List | VNAP | AUTO |
| Series Number | (0020,0011) | IS | - For Created Grayscale Softcopy Presentation State SOP Instances - Series Number equal to Series Number of | ALWAYS | AUTO |

|  |  |  | Referenced Image plus the number of Current Presentation State. - For Created Secondary Capture SOP Instances - Series Number is equal to Series Number of Referenced Image. - For Created Hardcopy SOP Instances - Series Number is equal to Series Number of Referenced Image - For Created Structured Report SOP Instances - Series number is generated if Structured Reports doesn’t exist for Study (otherwise existing series number of Structured Reports Series is used). - For Created PDF Documents SOP Instances - Series number is generated if PDF Document doesn’t exist for Study (otherwise existing Series number of Encapsulated PDF Documents Series is used). |  |  |
| --- | --- | --- | --- | --- | --- |

| Series Description | (0008,103E) | LO | - For Created Grayscale Softcopy Presentation State SOP Instances - Series Description of Referenced Image. - For Created Secondary Capture SOP Instances - Series Description of Existing Series. - For Created Hardcopy SOP Instances - empty string. - For Created Structured Report SOP Instances – empty string. - For Created PDF Documents SOP Instances – empty string. | VNAP | AUTO |
| --- | --- | --- | --- | --- | --- |
| Body Part Examined | (0018,0015) | CS | Body Part Examined | EMPTY | AUTO |
| Laterality | (0020,0060) | CS | Laterality | EMPTY | AUTO |

        - General Equipment Module

#### Table 9.1.1.6.9 - 1

General Equipment Module Attributes of Created SOP Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Manufacturer | (0008,0070) | LO | RamSoft Inc. | ALWAYS | AUTO |
| Institution Name | (0008,0080) | LO | - For Created Grayscale Softcopy Presentation State SOP Instances - Institution Name of Referenced Image. - For Created Secondary Capture SOP Instances - Institution Name of Existing Series. | ALWAYS | CONFIG |

|  |  |  | - For Created Hardcopy SOP Instances - Name of PowerServer. - For Created Structured Report SOP Instances – Institution Name of Related Study. For Created PDF Documents SOP Instances – Institution Name of Related Study. |  |  |
| --- | --- | --- | --- | --- | --- |
| Station Name | (0008,1010) | LO | - For Created Grayscale Softcopy Presentation State SOP Instances - Station Name of Referenced Image. - For Created Secondary Capture SOP Instances - Station Name of Existing Series. - For Created Hardcopy SOP Instances - Station Name of Existing Series. - For Created Structured Report SOP Instances – newer sent. For Created PDF Documents SOP Instances – newer sent. | VNAP | AUTO |
| Manufacturer Model Name | (0008,1090) | LO | RAMSOFT POWESERVER | ALWAYS | AUTO |
| Software Versions | (0018,1020) | LO | 6.0 | ALWAYS | AUTO |

#### Grayscale Softcopy Presentation State Modules (GSPS)

        - Presentation State

#### Table 9.1.1.7.1 - 1

Presentation State Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Instance Number | (0020,0013) | IS | Auto Generated (Number of Current Presentation State Instance of Referenced Image) | ALWAYS | AUTO |
| Presentation Label | (0070,0080) | CS | Auto Generated (Editable Name - Default – “ZOOM” + (Number of Current Presentation State Instance of Referenced Image) + Frame Number | ALWAYS | AUTO |
| Presentation Description | (0070,0081) | LO | Empty string | EMPTY | AUTO |
| Presentation Creation Date | (0070,0082) | DA | Auto Generated (Current Date) | ALWAYS | AUTO |
| Presentation Creation Time | (0070,0083) | TM | Auto Generated (Current Time) | ALWAYS | AUTO |
| Presentation Creator’s Name | (0070,0084) | PN | Auto Generated (User’s Name who created Presentation State) | ALWAYS | AUTO |
| Referenced Series Sequence | (0008,1115) | SQ | Auto Generated | ALWAYS | AUTO |
| > Series Instance UID | (0020,000E) | UI | Auto Generated | ALWAYS | AUTO |
| > Referred Image Sequence | (0008,1140) | SQ | Auto Generated | ALWAYS | AUTO |
| >> Referenced SOP Class UID | (0008,1150) | UI | Equal to SOP Class UID of Referenced Image | ALWAYS | AUTO |

| >> Referenced SOP Instance UID | (0008,1155) | UI | Equal to SOP Instance UID of Referenced Image | ALWAYS | AUTO |
| --- | --- | --- | --- | --- | --- |
| >> Referenced Frame Number | (0008,1160) | IS | Auto Generated (In case of Multifrtame Annotated all related frames of Referenced Image using VM delimiter) | ALWAYS | AUTO |

        - SOP Common

#### Table 9.1.1.7.2 - 1

SOP Common Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| SOP Class UID | (0008,0016) | UI | Grayscale Softcopy Presentation State Storage: “1.2.840.10008.5.1.4.1.1.11.1” | ALWAYS | AUTO |
| SOP Instance UID | (0008,0018) | UI | Generated automatically | ALWAYS | AUTO |
| Specific Character Set | (0008,0005) | UI | Equal to Presentation State of Referenced Image | ALWAYS | AUTO |
| Instance Creation Date | (0008,0012) | DA | Auto Generated (Date of Created Presentation State) | ALWAYS | AUTO |
| Instance Creation Time | (0008,0013) | TM | Auto Generated (Time of Created Presentation State) | ALWAYS | AUTO |
| Instance Creator UID | (0008,0014) | UI | 1.2.124.113540.0.0 | ALWAYS | AUTO |
| Timezone Offset From UTC | (0008,0201) | SH | Offset Timezone Offset From UTC of Referenced Image | ALWAYS | AUTO |

        - General Image

#### Table 9.1.1.7.3 - 1

General Image Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Content Date | (0008,0023) | DA | Auto Generated (Current Date) | ALWAYS | AUTO |
| Content Time | (0008,0033) | TM | Auto Generated (Current Time) | ALWAYS | AUTO |
| Acquisition Date | (0008,0022) | DA | Auto Generated (Current Date) | ALWAYS | AUTO |
| Acquisition Time | (0008,0032) | TM | Auto Generated (Current Time) | ALWAYS | AUTO |
| Acquisition DateTime | (0008,002A) | DT | Auto Generated (Current DateTime) | ALWAYS | AUTO |
| Acquisition Number | (0020,0012) | IS | -1 | ALWAYS | AUTO |

        - Display Shutter Module

#### Table 9.1.1.7.4 - 1

Display Shutter Module Attributes of Created GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Shutter Shape | (0018,1600) | CS | Enumerated Values: RECTANGULAR CIRCULAR POLYGONAL | ALWAYS | AUTO |
| Shutter Left Vertical Edge | (0018,1602) | IS | Sent if one value of Shutter Shape is rectangular | ALWAYS | AUTO |
| Shutter Right Vertical Edge | (0018,1604) | IS | Sent if one value of Shutter Shape is rectangular | ALWAYS | AUTO |
| Shutter Upper Horizontal Edge | (0018,1606) | IS | Sent if one value of Shutter Shape is rectangular | ALWAYS | AUTO |
| Shutter Lower Horizontal Edge | (0018,1608) | IS | Sent if one value of Shutter Shape is rectangular | ALWAYS | AUTO |
| Center of Circular Shutter | (0018,1610) | IS | Sent if one value of Shutter Shape is rectangular | ALWAYS | AUTO |

|  |  |  | (2 values with VM delimiter) |  |  |
| --- | --- | --- | --- | --- | --- |
| Radius of Circular Shutter | (0018,1612) | IS | Sent if one value of Shutter Shape is rectangular (2 – 2n values with VM delimiter) | ALWAYS | AUTO |
| Vertices of the Polygonal Shutter | (0018,1620) | IS | Sent if one value of Shutter Shape is Polygonal | ALWAYS | AUTO |
| Shutter Presentation Value | (0018,1622) | US | Auto Generated if Shutter is present | ALWAYS | AUTO |

        - Displayed Area Module

#### Table 9.1.1.7.5 - 1

Displayed Area Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Displayed Area Selection Sequence | (0070,005A) | SQ | Auto Generated | ALWAYS | AUTO |
| > Referenced Image Sequence | (0008,1140) | SQ | Auto Generated | ALWAYS | AUTO |
| >> Referenced SOP Class UID | (0008,1150) | UI | Equal to SOP Class UID of Referenced Image | ALWAYS | AUTO |
| >> Referenced SOP Instance UID | (0008,1155) | UI | Equal to SOP Instance UID of Referenced Image | ALWAYS | AUTO |
| >> Referenced Frame Number | (0008,1160) | IS | Frame Number of Referenced Image | ALWAYS | AUTO |
| > Displayed Area Top Left Hand Corner | (0070,0052) | SL | (2 values with VM delimiter) | ALWAYS | AUTO |
| > Displayed Area Top Bottom Right Corner | (0070,0053) | SL | (2 values with VM delimiter) | ALWAYS | AUTO |
| > Presentation State Mode | (0070,0100) | CS | Enumerated Values: SCALE TO FIT TRUE SIZE MAGNIFY | ALWAYS | AUTO |
| > Presentation Pixel Spacing | (0070,0101) | DS | The values are identical to PixelSpacing of Referenced | ALWAYS | AUTO |

|  |  |  | Image (2 values with VM delimiter) |  |  |
| --- | --- | --- | --- | --- | --- |
| > Presentation Pixel Spacing | (0070,0102) | DS | Generated if Presentation Pixel Spacing is not present | VNAP | AUTO |
| > Presentation Pixel Magnification Ratio | (0070,0103) | FL | Auto Generated | ALWAYS | AUTO |
| RamSoft Arbitrary Rotation Identifier | (310F,0010) | Private Tag (LO) | Generated if Arbitrary Rotation is present | VNAP | AUTO |
| Ramsoft Arbitrary Rotation | (310F,1010) | Private Tag (US) | Arbitrary Rotation Angle | VNAP | AUTO |
| RamSoft Apply Auto Size | (3121,1020) | Private Tag (CS) | Enumerated values: YES = Apply NO = Not Apply | ALWAYS | AUTO |
| RamSoft Fit To Width Identification | (3181, 0010) | Private Tag (LO) | Auto Generated Private Creator | ALWAYS | AUTO |
| RamSoft Fit To Width | (3181, 1010) | Private Tag (CS) | FIT TO WITH | VNAP | AUTO |
| RamSoft Image Horizontal Flip Identifier | (3185, 0010) | Private Tag (LO) | Auto Generated Private Creator | ALWAYS | AUTO |
| RamSoft Image Horizontal Flip | (3185,1010) | Private Tag (CS) | Enumerated values: Y, N | ALWAYS | AUTO |
| RamSoft Image Rotation Identifier | (3187,0010) | Private Tag (LO) | Auto Generated Private Creator | ALWAYS | AUTO |
| RamSoft Image Horizontal Flip | (3187,1010) | Private Tag (US) | Enumerated values: 0, 90, 180, 270 | ALWAYS | AUTO |

Note: In case of Multi-Frame Image Referenced Image Sequence is generated for each frame

        - Graphic Annotation Module

#### Table 9.1.1.7.6 - 1

Graphic Annotation Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presen ce of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Graphic Annotation Sequence | (0070,0001) | SQ | Generated only if Graphic Annotations are present | VNAP | AUTO |
| > Referenced Image Sequence | (0008, 1140) | SQ | Generated only if Graphic Annotations are present | VNAP | AUTO |
| >> Referenced SOP Class UID | (0008,1150) | UI | Equal to SOP Class UID of Referenced Image | VNAP | AUTO |
| >> Referenced SOP Instance UID | (0008,1155) | UI | Equal to Instance UID of Referenced Image | VNAP | AUTO |
| >> Referenced Frame Number | (0008,1160) | IS | Frame Number of Referenced Image. Sent in case of MultiFrame Referenced Image | VNAP | AUTO |
| > Graphic Layer | (0070, 0002) | CS | Generated only if Graphic Annotations are present | VNAP | AUTO |
| > Text Object Sequence | (0070, 0008) | SQ | Sent when text objectsor ROI objects are present in the graphic annotation | VNAP | AUTO |
| >> Bounding Box Annotation Units | (0070, 0003) | CS | Enumerated Value: PIXEL | VNAP | AUTO |
| >> Unformatted Text Value | (0070, 0006) | ST | In case of Text Annotation - Text Value entered by User In case of ROI: - Area of ROI in cm square - Perimiter pf ROI in cm - Mean of ROI - Std. Deviation of ROI | VNAP | USER |
| >> Bounding Box Top Left Hand Corner | (0070,0010) | FL | 2 values with VM delimiter | VNAP | USER |

| >> Bounding Box Bottom Right Hand Corner | (0070,0011) | FL | 2 values with VM delimiter | VNAP | USER |
| --- | --- | --- | --- | --- | --- |
| >> Bounding Box Text Horizontal Justification | (0070,0012) | CS | Enumerated Value: LEFT | VNAP | USER |
| >> Text Annotation Identification | (3179,0010) | Private Tag (LO) | Private Creator. Generated if Text Annotation is present | VNAP | AUTO |
| >> Referenced Graphic Annotation InstanceUID | (3179,1010) | Private Tag (UI) | Private Tag. Generated if Text Annotation is present | VNAP | AUTO |
| >> Show Text Annotation | (3179,1020) | Private Tag (CS) | Enumerated values: YES, NO. | VNAP | USER |
| > Graphic Object Sequence | (0070,0009) | SQ | Sent when graphic objects are present in the graphic annotation | VNAP | AUTO |
| >> Graphic Annotation Units | (0070,0005) | CS | Enumerated Value: PIXEL | VNAP | AUTO |
| >> Graphic Dimensions | (0070,0020) | US | Enumerated Value: 2 | VNAP | AUTO |
| >> Number of Graphic Points | (0070,0021) | US | 2 - n | VNAP | AUTO |
| >> Graphic Data | (0070,0022) | FL | (2 – 2n values with VM delimiter) | VNAP | AUTO |
| >> Graphic Type | (0070,0023) | CS | Enumerated Values: RECTANGULAR CIRCULAR POLYGONAL | VNAP | AUTO |
| >> Graphic Filled | (0070,0024) | CS | Enumerated Values: Y, N | VNAP | AUTO |

Note: In case of Multi-Frame Image Referenced Image Sequence is generated for each frame

        - Graphic Layer Module

This conditional module is sent only if graph annotations are present in the Presentation State module.

#### Table 9.1.1.7.7 - 1

Graphic Layer Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Graphic Layer Sequence | (0070,0060) | SQ | Generated only if Graphic Annotations are present | VNAP | AUTO |
| > Graphic Layer | (0070,0002) | CS | Graphic Layer Description Support only “Layer 0” value | VNAP | AUTO |
| > Graphic Layer Order | (0070,0062) | IS | Graphic Layer Order Support only First Layer | VNAP | AUTO |

        - Modality LUT Module

This conditional module is sent only if graph annotations are present in the Presentation State module.

#### Table 9.1.1.7.8 - 1

Modality LUT Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Resacale Intercept | (0028,1052) | DS | Copied from referenced image object in case of MONOCHROME Image | ALWAYS | AUTO |
| Resacale Slope | (0028,1053) | DS | Copied from referenced image object in case of MONOCHROME Image | ALWAYS | AUTO |
| Resacale Type | (0028,1054) | LO | Copied from referenced image object if exists in case of MONOCHROME Image. Otherwise in case of MONOCHROME Image value: “US” (unspecified). | ALWAYS | AUTO |

        - Softcopy VOI LUT Module

#### Table 9.1.1.7.9 - 1

Softcopy VOI LUT Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| SoftCopy VOI LUT Sequence | (0028,3110) | SQ |  | ALWAYS | AUTO |
| > Referenced Image Sequence | (0008,1140) | SQ |  | ALWAYS | AUTO |
| >> Referenced SOP Class UID | (0008,1150) | UI | Equal to SOP Class UID of Referenced Image | ALWAYS | AUTO |
| >> Referenced SOP Instance UID | (0008,1155) | UI | Equal to SOP Instance UID of Referenced Image | ALWAYS | AUTO |
| >> Referenced Frame Number | (0008,1160) | IS | Frame Number of Referenced Image. Sent in case of Multi-Frame Referenced Image | ALWAYS | AUTO |
| Window Center | (0028,1050) | DS | User established value | ALWAYS | AUTO |
| Window Center | (0028,1051) | DS | User established value | ALWAYS | AUTO |

        - Softcopy Presentation LUT Module

#### Table 9.1.1.7.10 - 1

Softcopy Presentation LUT Module Attributes of GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Presentation LUT Shape | (2050, 0010) | CS | Enumerated Values: NORMAL, INVERSE | VNAP | AUTO |

#### Secondary Capture IOD

        - SC Equipment

#### Table 9.1.1.8.1 - 1

SC Equipment Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Conversion Type | (0008,0064) | CS | Enumerated Value: WSD | ALWAYS | AUTO |

        - General Image

#### Table 9.1.1.8.2 - 1

General Image Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Content Date | (0008,0023) | DA | Auto Generated (Current Date) | ALWAYS | AUTO |
| Content Time | (0008,0033) | TM | Auto Generated (Current Time) | ALWAYS | AUTO |
| Acquisition Date | (0008,0022) | DA | Auto Generated (Current Date) | ALWAYS | AUTO |
| Acquisition Time | (0008,0032) | TM | Auto Generated (Current Time) | ALWAYS | AUTO |
| Acquisition DateTime | (0008,002A) | DT | Auto Generated (Current DateTime) | ALWAYS | AUTO |
| Acquisition Number | (0020,0012) | IS | -1 | ALWAYS | AUTO |
| Instance Number | (0020,0013) | IS | Increrased Last Image Number in Series | ALWAYS | AUTO |
| Patient Orientation | (0020,0020) | CS | Auto Generated | EMPTY | AUTO |
| Frame of ReferenceUID | (0020,0052) | CS | Auto Generated | EMPTY | AUTO |
| Time zone Offset From UTC | (0008,0201) | SH | Equal to Time zone Offset of the Study. | ALWAYS | AUTO |
| Image Type | (0008,0008) | CS | Enumerated value: DERIVED\PRIMARY | ALWAYS | AUTO |

        - Image Pixel

#### Table 9.1.1.8.39.1.1.8.2 - 1

Image Pixel Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Transfer Syntax | (0002,0010) | UI | JPEG 2000 LossLess Only 1.2.840.10008.1.2.4.90 | ALWAYS | AUTO |
| Samples per Pixel | (0028,0002) | US | Grayscale Image: 1 Color Image: 3 | ALWAYS | AUTO |
| Photometric Interpretation | (0028,0004) | CS | Grayscale Image: MONOCHROME2 Color Image: YBR_RCT | ALWAYS | AUTO |
| Number of Frames | (0028,0008) | IS | 1 –n | ALWAYS | AUTO |
| Rows | (0028,0010) | US | Number of rows of imported image | ALWAYS | AUTO |
| Columns | (0028,0011) | US | Number of columns of imported image | ALWAYS | AUTO |
| Bits Allocated | (0028,0100) | US | Enumerated Value: 8 | ALWAYS | AUTO |
| Bits Stored | (0028,0101) | US | Enumerated Value: 8 | ALWAYS | AUTO |
| High Bit | (0028,0102) | US | Enumerated Value: 7 | ALWAYS | AUTO |
| Pixel Representation | (0028,0103) | US | Enumerated Value: 0 | ALWAYS | AUTO |
| Planar Configuration | (0028,0006) | US | Enumerated Value: 1 | ALWAYS | AUTO |
| Lossy Image Compression | (0028,2110) | CS | Enumerated Value: 00 | ALWAYS | AUTO |
| Pixel Data | (7FE0,0010) | OB | Always generated as OB | ALWAYS | AUTO |

        - SOP Common Module

#### Table 9.1.1.8.49.1.1.8.2 - 1

SOP Common Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presenc e of Value | Source |
| --- | --- | --- | --- | --- | --- |
| SOP Class UID | (0008,0016) | UI | Secondary Capture Image Storage: 1.2.840.10008.5.1.4.1.1.7 | ALWAYS | AUTO |
| SOP Instance UID | (0008,0018) | UI | Generated automatically | ALWAYS | AUTO |

| Specific Character Set | (0008,0005) | CS | UTF8 Character Set ISO_IR 192 | ALWAYS | AUTO |
| --- | --- | --- | --- | --- | --- |
| Instance Creation Date | (0008,0012) | DA | Auto Generated (Date of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creation Time | (0008,0013) | TM | Auto Generated (Time of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creator UID | (0008,0014) | UI | 1.2.124.113540.0.0 | ALWAYS | AUTO |
| Timezone Offset From UTC | (0008,0201) | SH | Offset Timezone Offset From UTC of Study | ALWAYS | AUTO |

        - Modality LUT Module

#### Table 9.1.1.8.59.1.1.8.2 - 1

Modality LUT Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Resacale Intercept | (0028,1052) | DS | 0.000000 in case of Grayscale Image | ALWAYS | AUTO |
| Resacale Slope | (0028,1053) | DS | 1.000000 in case of Grayscale Image | ALWAYS | AUTO |
| Resacale Type | (0028,1054) | DS | US in case of Grayscale Image | ALWAYS | AUTO |

        - Softcopy VOI LUT Module

Table 9.1.1.8.6 - 1

Softcopy VOI LUT Module Attributes of Secondary Capture Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Window Center | (0028,1050) | DS | In case of Grayscale Image: 128 | ALWAYS | AUTO |
| Window Width | (0028,1051) | DS | In case of Grayscale Image: 256 | ALWAYS | AUTO |

#### Hardcopy IOD

        - SC Equipment

#### Table 9.1.1.9.1 - 1

Hardcopy Equipment Module Attributes of Hardcopy Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Conversion Type | (0008,0064) | CS | Enumerated Value: DF | ALWAYS | AUTO |

        - General Image

#### Table 9.1.1.9.2 - 1

General Image Module Attributes of Hardcopy Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Content Date | (0008,0023) | DA | Auto Generated (Date of Created Hardcopy Image) | ALWAYS | AUTO |
| Content Time | (0008,0033) | TM | Auto Generated (Time of Created Hardcopy Image) | ALWAYS | AUTO |
| Acquisition Date | (0008,0022) | DA | Auto Generated (Date of Created Hardcopy Image) | ALWAYS | AUTO |
| Acquisition Time | (0008,0032) | TM | Auto Generated (Time of Created Hardcopy Image) | ALWAYS | AUTO |
| Acquisition DateTime | (0008,002A) | DT | Auto Generated (Date and Time of Created Hardcopy Image) | ALWAYS | AUTO |
| Acquisition Number | (0020,0012) | IS | -1 | ALWAYS | AUTO |
| Instance Number | (0020,0013) | IS | 1 | ALWAYS | AUTO |
| Patient Orientation | (0020,0020) | CS | Auto Generated | EMPTY | AUTO |
| Frame of ReferenceUID | (0020,0052) | CS | Auto Generated | EMPTY | AUTO |
| Image Type | (0008,0008) | CS | Enumerated value: DERIVED\PRIMARY | ALWAYS | AUTO |

        - Image Pixel

#### Table 9.1.1.9.39.1.1.8.2 - 1

Image Pixel Module Attributes of Hardcopy Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Transfer Syntax | (0002,0010) | UI | JPEG LossLess, Non-hierarchical, 1st Order Prediction 1.2.840.10008.1.2.4.70 | ALWAYS | AUTO |
| Samples per Pixel | (0028,0002) | US | Grayscale Image: 1 Color Image: 3 | ALWAYS | AUTO |
| Photometric Interpretation | (0028,0004) | CS | Grayscale Image: MONOCHROME2 Color Image: RGB | ALWAYS | AUTO |
| Number of Frames | (0028,0008) | IS | 1 | ALWAYS | AUTO |
| Rows | (0028,0010) | US | Number of rows of scanned image | ALWAYS | AUTO |
| Columns | (0028,0011) | US | Number of columns of scanned image | ALWAYS | AUTO |
| Bits Allocated | (0028,0100) | US | Enumerated Value: 8 | ALWAYS | AUTO |
| Bits Stored | (0028,0101) | US | Enumerated Value: 8 | ALWAYS | AUTO |
| High Bit | (0028,0102) | US | Enumerated Value: 7 | ALWAYS | AUTO |
| Pixel Representation | (0028,0103) | US | Enumerated Value: 0 | ALWAYS | AUTO |
| Planar Configuration | (0028,0006) | US | Enumerated Value: 1 | ALWAYS | AUTO |
| Imager Pixel Spacing | (0018,1164) | DS | Default Values in mm: 0.1693333\0.1693333 | ALWAYS | CONFIG |
| Burned In Annotation | (0028,0301) | CS | Enumerated Value: YES | ALWAYS | AUTO |
| Lossy Image Compression | (0028,2110) | CS | Enumerated Value: 00 | ALWAYS | AUTO |
| Pixel Data | (7FE0,0010) | OB | Always generated as OB | ALWAYS | AUTO |

        - SOP Common Module

#### Table 9.1.1.9.49.1.1.8.2 - 1

SOP Common Module Attributes of Hardcopy Instances

| Attribute Name | Tag | VR | Value | Presenc e of Value | Source |
| --- | --- | --- | --- | --- | --- |
| SOP Class UID | (0008,0016) | UI | Multiframe True Color Secondary Capture Image Storage: 1.2.840.10008.5.1.4.1.1.7.4 | ALWAYS | AUTO |
| SOP Instance UID | (0008,0018) | UI | Generated automatically | ALWAYS | AUTO |
| Specific Character Set | (0008,0005) | CS | UTF8 Character Set ISO_IR 192 | ALWAYS | AUTO |
| Instance Creation Date | (0008,0012) | DA | Auto Generated (Date of Created Hardcopy Image) | ALWAYS | AUTO |
| Instance Creation Time | (0008,0013) | TM | Auto Generated (Time of Created Hardcopy Image) | ALWAYS | AUTO |
| Instance Creator UID | (0008,0014) | UI | 1.2.124.113540.0.0 | ALWAYS | AUTO |
| Timezone Offset From UTC | (0008,0201) | SH | Offset Timezone Offset From UTC of Study | ALWAYS | AUTO |

        - Modality LUT Module

#### Table 9.1.1.9.59.1.1.8.2 - 1

Modality LUT Module Attributes of Hardcopy Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Resacale Intercept | (0028,1052) | DS | 0.000000 in case of Grayscale Image | ALWAYS | AUTO |
| Resacale Slope | (0028,1053) | DS | 1.000000 in case of Grayscale Image | ALWAYS | AUTO |
| Resacale Type | (0028,1054) | DS | US in case of Grayscale Image | ALWAYS | AUTO |

        - Softcopy VOI LUT Module

#### Table 9.1.1.9.6 - 1

Softcopy VOI LUT Module Attributes of Created GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Window Center | (0028,1050) | DS | In case of Grayscale Image: 128 | ALWAYS | AUTO |
| Window Width | (0028,1051) | DS | In case of Grayscale Image: 256 | ALWAYS | AUTO |

        - Softcopy Presentation LUT Module

Table 9.1.1.9.7 - 1

Softcopy Presentation LUT Module Attributes of Created GSPS Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Presentation LUT Shape | (2050,0010) | CS | In case Grayscale Image Enumerated Values: NORMAL | ALWAYS | AUTO |

#### Encapsulated PDF Document IOD

        - SC Equipment

#### Table 9.1.1.10.1 - 1

SC Equipment Module Attributes of Encapsulated PDF Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Conversion Type | (0008,0064) | CS | Enumerated Value: WSD | ALWAYS | AUTO |

        - General Image

Table Error! Reference source not found. - 1

#### General Image Module Attributes of Encapsulated PDF Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Content Date | (0008,0023) | DA | Auto Generated (The date the document content creation was started) | ALWAYS | AUTO |
| Content Time | (0008,0033) | TM | Auto Generated (The time the document content creation was started) | ALWAYS | AUTO |
| Acquisition Date | (0008,0022) | DA | Auto Generated (The date the document content creation was started) | ALWAYS | AUTO |
| Acquisition Time | (0008,0032) | TM | Auto Generated (The time the document content creation was started. | ALWAYS | AUTO |
| Acquisition DateTime | (0008,002A) | DT | Auto Generated (The date and time the document content creation was started) | ALWAYS | AUTO |
| Acquisition Number | (0020,0012) | IS | -1 | ALWAYS | AUTO |
| Instance Number | (0020,0013) | IS | Number of Document in PDF Document Series | ALWAYS | AUTO |

        - Encapsulated Document

#### Table 9.1.1.10.3 - 1

Encapsulated Document Module Attributes of Encapsulated PDF Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Concept Name Code Sequence | (0040,A043) | SQ | A coded representation of the document title. | ALWAYS | AUTO |
| >> Code Value | (0008, 0100) | SH | DT.01 | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008, 0104) | LO | Enumerated Values: Study Level Patient Level | ALWAYS | AUTO |
| Burned In Annotation | (0028, 0301) | DA | NO | ALWAYS | AUTO |
| Document Title | (0042, 0010) | ST | Identifies the Title of the Document. For Example (PRELIMINARY, FINAL) | ALWAYS | AUTO |
| MIME Type of Encapsulated Document | (0042, 0012) | LO | application/pdf | ALWAYS | AUTO |
| Encapsulated Document | (0042, 0011) | OB | Encapsulated Document stream, containing a document encoded according to the MIME Type | ALWAYS | AUTO |
| Content Creator Name | (0070,0084) | PN | 99RAMSOFT | ALWAYS | AUTO |
| RamSoft Document Type Identifier | (311B, 0010) | Private Tag (LO) | RamSoft Document Type Identifier (Private Creator) | ALWAYS | AUTO |
| Document Type | (311B, 1010) | Private Tag (LO) | Document type of Encapsulated PDF Document. See Enumerated Values of Private Tag (311B, 1010) Paragraph 9.2 | ALWAYS | AUTO |

        - SOP Common Module

Table 9.1.1.10.4 - 1

SOP Common Module Attributes of of Encapsulated PDF Document Instances

| Attribute Name | Tag | VR | Value | Presenc e of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Transfer Syntax UID | (0002,0010) | UI | Deflated Little Endian Explicit 1.2.840.10008.1.2.1.99 | ALWAYS | AUTO |
| SOP Class UID | (0008, 0016) | UI | Encapsulated PDF Storage: 1.2.840.10008.5.1.4.1.1.104.1 | ALWAYS | AUTO |
| SOP Instance UID | (0008, 0018) | UI | Generated automatically | ALWAYS | AUTO |
| Specific Character Set | (0008, 0005) | CS | UTF8 Character Set ISO_IR 192 | ALWAYS | AUTO |
| Instance Creation Date | (0008, 0012) | DA | Auto Generated (Date of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creation Time | (0008, 0013) | TM | Auto Generated (Time of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creator UID | (0008, 0014) | UI | 1.2.124.113540.0.0 | ALWAYS | AUTO |
| Timezone Offset From UTC | (0008, 0201) | SH | Offset Timezone Offset From UTC of Study | ALWAYS | AUTO |
| Coding Scheme Identification Sequence | (0008, 0110) | SQ | Sequence of items that map values of Coding Scheme Designator (0008,0102) to the local coding scheme | ALWAYS | AUTO |
| > Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| > Coding Scheme Version | (0008, 0103) | SH | 6.0 | ALWAYS | AUTO |
| > Coding Scheme UID | (0008, 010C) | UI | 1.2.124.113540.1.1.6 | ALWAYS | AUTO |
| > Coding Scheme Name | (0008, 0115) | ST | RAMSOFT Coding Scheme | ALWAYS | AUTO |
| > Responsible Organization | (0010, 2299) | LO | RamSoft Inc. | ALWAYS | AUTO |

#### Structured Report Document IOD

        - General Image

#### Table 9.1.1.11.1 - 1

General Image Module Attributes of Structured Report Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Content Date | (0008,0023) | DA | Auto Generated (The date the document content creation was started) | ALWAYS | AUTO |
| Content Time | (0008,0033) | TM | Auto Generated (The time the document content creation was started) | ALWAYS | AUTO |
| Acquisition Date | (0008,0022) | DA | Auto Generated (The date the document content creation was started) | ALWAYS | AUTO |
| Acquisition Time | (0008,0032) | TM | Auto Generated (The time the document content creation was started. | ALWAYS | AUTO |
| Acquisition DateTime | (0008,002A) | DT | Auto Generated (The date and time the document content creation was started) | ALWAYS | AUTO |
| Acquisition Number | (0020,0012) | IS | -1 | ALWAYS | AUTO |
| Instance Number | (0020,0013) | IS | Number of Document in SR Document Series | ALWAYS | AUTO |

        - Structured Document General

#### Table 9.1.1.11.2 - 1

Structured Document General Module Attributes of Structured Report Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Completion Flag | (0040,A491) | CS | Enumerad values: PARTIAL COMPLETED | ALWAYS | AUTO |
| Verification Flag | (0040,A493) | CS | Enumerad values: UNVERIFIED VERIFIED | ALWAYS | AUTO |

| Verification Observer Sequence | (0040,A073) | SQ | Sent if report was verified | VNAP | AUTO |
| --- | --- | --- | --- | --- | --- |
| > Verifying Organization | (0040,A027) | LO | Institution Name if the Study | ALWAYS | AUTO |
| > Verification Date Time | (0040,A030) | DT | Auto Generated (The date and time the document content creation was started) | ALWAYS | AUTO |
| > Verifying Observer Name | (0040,A075) | PN | Full Name of User who performed report | ALWAYS | AUTO |
| > Verification Observer Identification Code Sequence | (0040,A088) | SQ | Sent if User exists in RamSoft PACS | VNAP | AUTO |
| >> Code Value | (0008,0100) | SH | Identification Code of User in RamSoft PACS | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008,0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008,0104) | LO | Enumerated value: Observer | ALWAYS | AUTO |
| Predecessor Documents Sequence | (0040,A360) | SQ | Generated if Document was verified (signed) | VNAP | AUTO |
| > Study Instance UID | (0020,000D) | UI | Study Instance UID of examinated study | ALWAYS | AUTO |
| > Referenced Series Sequence | (0008,1155) | SQ | Generated if Document was verified (signed) | VNAP | AUTO |
| >> Series Instance UID | (0020,000E) | UI | Series Instance UID of Structured Report Document Series | ALWAYS | AUTO |
| >> Referenced SOP Sequence | (0008,1199) | SQ | Generated if Document was verified (signed) | VNAP | AUTO |
| >>> Referenced SOP ClassUID | (0008,1150) | UI | Value: BasicTextSRStorage | ALWAYS | AUTO |
| >>> Referenced SOP Instance UID | (0008,1155) | UI | SOP Instance UID of Predecessor unverified Structured Report | ALWAYS | AUTO |
| Performed Procedure Code Sequence | (0040,A372) | SQ | Empty sequence sent | EMPTY | AUTO |

        - Structured Document Content (Basic Text)

#### Table 9.1.1.11.3 - 1

Structured Document Content (Text Content) Module Attributes of Structured Report Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Value Type | (0040,A040) | CS | Enumerated Value: CONTAINER | ALWAYS | AUTO |
| Concept Name Code Sequence | (0040,A043) | SQ | A coded representation of the document title | ALWAYS | AUTO |
| >> Code Value | (0008,0100) | SH | DT.01 | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008,0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008,0104) | LO | Identifies the Title of the Document. For Example (PRELIMINARY, FINAL) | ALWAYS | AUTO |
| Continuity Of Content | (0040,A050) | CS | Enumerated Value: SEPARATE | ALWAYS | AUTO |
| Content Sequence | (0040,A730) | SQ |  | ALWAYS | USER |
| > Relationship Type | (0040,A010) | CS | CONTAINS | ALWAYS | USER |
| > ValueType | (0040,A040) | CS | TEXT | ALWAYS | USER |
| > Concept Name Code Sequence | (0040,A043) | SQ | A coded representation of the document header | ALWAYS | AUTO |
| >> Code Value | (0008, 0100) | SH | UT_n (n – Identification of the Documnent Header) | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008, 0104) | LO | The header of the Document | ALWAYS | AUTO |
| > Text Value | (0040,A160) | UT | Text editedby the User | USER | AUTO |
| RamSoft Document Type Identifier | (311B, 0010) | Private Tag (LO) | RamSoft Document Type Identifier (Private Creator) | ALWAYS | AUTO |
| Document Type | (311B, 1010) | Private Tag (LO) | Document type of Strucured Report Document. See Enumerated Values of Private Tag (311B, 1010) Paragraph 9.2 | ALWAYS | AUTO |

        - Structured Document Content (Binary Content)

#### Table 9.1.1.11.4 - 1

Structured Document Content Module Attributes of Structured Report Document Instances

| Attribute Name | Tag | VR | Value | Presence of Value | Source |
| --- | --- | --- | --- | --- | --- |
| Value Type | (0040,A040) | CS | Enumerated Value: CONTAINER | ALWAYS | AUTO |
| Concept Name Code Sequence | (0040,A043) | SQ | A coded representation of the document title. | ALWAYS | AUTO |
| >> Code Value | (0008, 0100) | SH | DT.01 | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008, 0104) | LO | Identifies the Title of the Document. For Example (PRELIMINARY, FINAL) | ALWAYS | AUTO |
| Continuity Of Content | (0040,A050) | CS | Enumerated Value: SEPARATE | ALWAYS | AUTO |
| Content Sequence | (0040,A730) | SQ |  | ALWAYS | USER |
| > Relationship Type | (0040,A010) | CS | CONTAINS | ALWAYS | USER |
| > ValueType | (0040,A040) | CS | TEXT | ALWAYS | USER |
| > Concept Name Code Sequence | (0040,A043) | SQ | A coded representation of the document header | ALWAYS | AUTO |
| >> Code Value | (0008, 0100) | SH | UT_n (n – Identification of the Documnent Header) | ALWAYS | AUTO |
| >> Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| >> Code Meaning | (0008, 0104) | LO | Enumerated Value: Report | ALWAYS | AUTO |
| > Text Value | (0040,A160) | UT | The Text translated from Binary Document | USER | AUTO |
| RamSoft File Kind Identifier | (3111, 0010) | Private Tag (LO) | RamSoft File Kind Identifier (Private Creator) | ALWAYS | AUTO |
| File Kind | (3111, 1010) | Private Tag (CS) | RamSoft File Kind of Strucured Report | ALWAYS | AUTO |

|  |  |  | Document. See Enumerated Values of Private Tag (3111, 1010) Paragraph 9.2 |  |  |
| --- | --- | --- | --- | --- | --- |
| RamSoft Custom Report Identifier | (3113, 0010) | LO | RamSoft Custom Report Identifier (Private Creator) | ALWAYS | AUTO |
| Custom Report Data | (3113, 1010) | OB | Binary Document stream, containing a document encoded according to Document Kind | ALWAYS | AUTO |
| Custom Report Data Size | (3113, 1020) | UL | Size of Custom Report Data in bytes | ALWAYS | AUTO |
| Custom Report Template Data | (3113, 1030) | OB | Custom report Word template data. Generated in case of Word Template | VNAP | AUTO |
| Custom Report Template Data Size | (3113, 1040) | UL | Size in bytes of custom report Word template in bytes. Generated in case of Word Template. | VNAP | AUTO |
| RamSoft Document Type Identifier | (311B, 0010) | Private Tag (LO) | RamSoft Document Type Identifier (Private Creator) | ALWAYS | AUTO |
| Document Type | (311B, 1010) | Private Tag (LO) | Document type of Strucured Report Document. See Enumerated Values of Private Tag (311B, 0010) Paragraph 9.2 | ALWAYS | AUTO |

        - SOP Common Module

Table 9.1.1.11.5 - 1

SOP Common Module Attributes of of Structured Report Document Instances

| Attribute Name | Tag | VR | Value | Presenc e of Value | Source |
| --- | --- | --- | --- | --- | --- |

| Transfer Syntax UID | (0002,0010) | UI | Deflated Little Endian Explicit 1.2.840.10008.1.2.1.99 | ALWAYS | AUTO |
| --- | --- | --- | --- | --- | --- |
| SOP Class UID | (0008, 0016) | UI | Basic Text SR Storage: 1.2.840.10008.5.1.4.1.1.88.11 Enchanced SR Storage: 1.2.840.10008.5.1.4.1.1.88.22 Comprehensive SR Storage: 1.2.840.10008.5.1.4.1.1.88.33 | ALWAYS | AUTO |
| SOP Instance UID | (0008, 0018) | UI | Generated automatically | ALWAYS | AUTO |
| Specific Character Set | (0008, 0005) | CS | UTF8 Character Set ISO_IR 192 | ALWAYS | AUTO |
| Instance Creation Date | (0008, 0012) | DA | Auto Generated (Date of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creation Time | (0008, 0013) | TM | Auto Generated (Time of Created Secondary Capture Image) | ALWAYS | AUTO |
| Instance Creator UID | (0008, 0014) | UI | 1.2.124.113540.0.0 | ALWAYS | AUTO |
| Timezone Offset From UTC | (0008, 0201) | SH | Offset Timezone Offset From UTC of Study | ALWAYS | AUTO |
| Coding Scheme Identification Sequence | (0008, 0110) | SQ | Sequence of items that map values of Coding Scheme Designator (0008,0102) to the local coding scheme | ALWAYS | AUTO |
| > Coding Scheme Designator | (0008, 0102) | SH | 99RAMSOFT | ALWAYS | AUTO |
| > Coding Scheme Version | (0008, 0103) | SH | 6.0 | ALWAYS | AUTO |
| > Coding Scheme UID | (0008, 010C) | UI | 1.2.124.113540.1.1.6 | ALWAYS | AUTO |
| > Coding Scheme Name | (0008, 0115) | ST | RAMSOFT Coding Scheme | ALWAYS | AUTO |
| > Responsible Organization | (0010, 2299) | LO | RamSoft Inc. | ALWAYS | AUTO |

## Usage of Attributes From Received IODs

DICOM acquired images are stored AS IS by RamSoft PACS. Images that do not contain the mandatory fields for RamSoft PACS are rejected.

Table 9.1.2 -1 Image IOD Specifications

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Patient's Name | (0010,0010) | 2 | Patient | This field is mandatory for RamSoft PACS. At a minimum, the last name must be present. |
| Patient ID | (0010,0020) | 2 | Patient | If field is not present or blank, it will be auto generated. |
| Issuer Of PatientID | (0010,0021) | 3 | Patient | If field is not present or blank, it will be auto generated. |
| Patient’s Birth Date | (0010,0030) | 3 | Patient | Date of birth of the named patient |
| Patient’s Sex | (0010,0040) | 3 | Patient | Sex of the named patient. Enumerated Values: M = male F = female O = other |
| Patient's Insurance Plan Code Sequence | (0010,0050) | 3 | Patient | A sequence that conveys the patient's insurance plan. Zero or more Items is included in this Sequence. |
| > Code Value | (0008,0100 | 3 | Patient | Values. For Example, “12345” (Account Number) |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Patient’s Primary Language Code Sequence | (0010,0101) | 3 | Patient | The languages that used to communicate with the patient. |
| > Code Value | (0008,0100 | 3 | Patient | Values. If this is absent, the language is set to “DECLINED”. |
| Other Patient IDs | (0010,1000 | 3 | Patient | SSN or Health Card Number |
| Patient’s Size | (0010,1020) | 3 | Patient | Patient height or length in meters |
| Patients’ Weight | (0010,1030) | 3 | Patient | Patient weight in kilograms |
| Patient’s Address | (0010,1040) | 3 | Patient | The format of address is following: Street Address^City^State^Zip For Example: 10808 FOOTHILL BLVD^RANCHO CUCAMONGA^CA^91730 |
| Country of Residence | (0010,2150) | 3 | Patient | Country code in which patient currently resides. We store only the first 2 characters as we are expecting a 2 character country code. |
| Patient’s Telephone Numbers | (0010,2154) | 3 | Patient | Telephone numbers at which the patient can be reached. (Only Digits are exported). |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Ethnic Group | (0010,2160) | 3 | Patient | Ethnicity of patient. Enumerated values: 2135-2 (HISPANIC OR LATINO) 2186-5 (NOT HISPANIC OR LATINO) UNK (DECLINED). |
| Smoking Status | (0010,21A0) | 3 | Patient | Indicates whether patient smokes. Enumerated values: EVERYDAY SMOKER SOMEDAY SMOKER FORMER SMOKER NEVER SMOKER UNKNOWN SMOKER UNKNOWN |
| Additional Patient History | (0010,21B0) | 3 | Patient | Additional information about the patient's medical history |
| Allergies | (0010,2110) | 3 | Patient | The format is following: ALLERGY NAME\ALLERGY NOTES\SEVERITY Enumerated Values for Severity are following: “MI” – Mild ‘MO” – Moderate “SV” – Severe For Example: PENICILLIN\PRODUCES HIVES\SV |
| Admission ID | (0038, 0010) | 3 | Visit | Visit ID |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Route of Admissions | (0038, 0016) | 3 | Visit | Mode of admission. Enumerated values: E = EMERGENCY I = INPATIENT O = OUTPATIENT P = PREADMIT R = RECURRING PATIENT B = OBSTETRICS |
| Admitting Date | (0038, 0020) | 3 | Visit | Date patient visit began |
| Admitting Time | (0038, 0021) | 3 | Visit | Time patient visit began |
| Current Patient Location | (0038,0300) | 3 | Visit | The current department of the patient |
| Patient’s Institution Residence | 0038,0400) | 3 | Visit | The current room where patient resides |
| Visit Comments | (0038, 4000) | 3 | Visit | User-defined comments about the visit |
| Study Instance UID | 0020,000D) | 1 | Study | This field is mandatory for RamSoft PACS. |
| Accession Number | (0008,0050) | 2 | Study | This field is essential for RIS connectivity, but auto-generated if blank. |
| StudyID | (0020,0010 | 1 | Study | Identification of the Study. Auto-generated if blank. |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| StudyStatusID | (0032,000A) | 2 | Study | Status values are customized for each site. The following lists only our default values: ORDERED SCHEDULED CONFIRMED ARRIVED CANCELLED READYFORSCAN STARTED EXAMCOMPLETED DISCONTINUED INPROGRESS COMPLETED HOLD REJECTED VERIFIED TO BE AMENDED DICTATED TRANSCRIBED SIGNED PRIOR |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Study Priority ID | (0032,000C) | 3 | Study | Priority statuses are customized for each site. The following lists only our default values: LOW (ROUTINE) MED (MEDIUM) HIGH (HIGH) STAT (STAT) CRITTEST (CRITTEST) STAT (CRITFIND) |
| Study Description | (0008,1030) | 3 | Study | Study Description |
| Requested Procedure ID | (0040,1001) | 3 | Study | Requested Procedure ID |
| Institution Name | (0008,0080) | 3 | Study | Institution Name. |
| Department Name | (0008,1040) | 3 | Study | Department Name. |
| Procedure Code Sequence | (0008,1032) | 3 | Study | Zero or more Items is included in this Sequence. |
| > CodeValue | (0008,0100) | 3 | Study | Procedure Code (For Example “70100”). |
| > Coding Scheme Designator | (0008,0102) | 3 | Study | Coding Scheme Designator (For Example “DCM”) |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| > Code Meaning | (0008,0104) | 3 | Study | Procedure Description (For Example “XRAY JAW < 4 VIEWS” ) |
| Admitting Diagnoses Code Sequence | (0008,1084) | 3 | Study | Zero or more Items is included in this Sequence. |
| > CodeValue | (0008,0100) | 3 | Study | Diagnosis Code (For Example “830.1”) |
| > Coding Scheme Designator | (0008,0102) | 3 | Study | Coding Scheme Designator (For Example “DCM”) |
| > Code Meaning | (0008,0104) | 3 | Study | Diagnosis Description (For Example “DISLOCATION OF JAW; OPEN DISLOCATION”) |
| Reason for Study | (0032,1030) | 3 | Study | Symptoms |
| Study Comments | (0032,4000) | 3 | Study | Study Comments |
| Referring Physician’s Name | (0008,0090) | 3 | Study | Referring Physician |
| Performing Physician’s Name | (0008,1050) | 3 | Study | Performing Physician |
| Name Of Physician Reading Study | (0008,1060) | 3 | Study | Reading Physician |
| Operators Name | (0008,1070) | 3 | Study | Performing Technologist |

| Entity Name | Tag | Type | Module Name | Details |  |
| --- | --- | --- | --- | --- | --- |
| Physicians Of Record | (0008,1048) | 3 Value Multi plicity (1-n) | Study | Consulting Physicians. For Example 2 Consulting Physicians: (JANG^ANTHONY\PATIL^ANIL). |  |
| Interpretation Transcriber | (4008,010a) | 3 | Study | Transcriptionist |  |
| Series Instance UID | (0020,000E) | 1 | General Series | This field is mandatory for RamSoft PACS. |  |
| Series Number | (0020,0011) | 2 | General Series | The first series of a study is assigned the number 1 for acquired and imported images. This acquisition number is incremented for each successive image. |  |
| Modalty | (0008,0060) | 1 | General Series | Modalty. For Example (“CT”) |  |
| Laterality | (0020,0062) | 2C | General Series | Laterality. For Exampe (“L”) |  |
| Body Part Examined | (0018,0015) | 3 | General Series | Body Part Examined. For Example (“CHEST”) |  |
| Specific Character Set | (0008,0005) | 3 | SOP Common | Specific Character Set. For Example (ISO_IR 100) |  |
| Manufacturer | (0008,0070) | 2 | General Equipment | Manufacturer. For Example (“GE MEDICAL SYSTEM”). |  |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Manufacturer's Model Name | (0008,1090) | 3 | General Equipment | Manufacturer's Model Name. For Example “Optima MR450w”. |
| Software Versions | (0018,1020) | 3 | General Equipment | Software Versions. For Example (“23\LX\MR Software release:DV22.0_V02_1122.a). |
| Image Type | (0008,0008) | 3 | General Image and US Image | Image Type. For Example [DERIVED\PRIMARY]. |
| Acquisition Number | (0020,0012) | 3 | General Image | Acquisition Number. |
| Time zone Offset From UTC | (0008,0201) | 3 | General Image | Contains the offset from UTC to the time zone for all DA and TM Attributes present in this SOP Instance, and for all DT Attributes present in this SOP Instance that do not contain an explicitly encoded time zone offset. |
| Referenced Image Sequence | (0008,1140) | 3 | General Image | This is set and used to display scout thumbnails on CT/MR images. |
| >Referenced SOP Class UID | (0008,1150) | 1C | General Image | Set and used if Referenced Image Sequence is present. |
| >Referenced SOP Instance UID | (0008,1155) | 1C | General Image | Set and used if Referenced Image Sequence is present. |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Derivation Description | (0008,2111) | 3 | General Image | If JPEG lossy compression is used to save the image, this is set to "JPEG Lossy N:1" to indicate the lossy compression ratio used. This field is displayed as an overlay on the image to indicate the lossy compression ratio of images. |
| Burned in Annotation | (0028,0301) | 3 | General Image | For Example “YES”. |
| Lossy Image Compression | (0028,2110) | 3 | General Image | Set to “01” if the image has been subjected to lossy image compression. |
| Lossy Image Compression Ratio | (0028,2112) | 3 | General Image | Set to the approximate lossy compression ratio that has been applied to the image e.g. 30 for 30:1 compression. |
| Pixel Spacing | (0028,0030) | 1 | Image Plane | Used to establish the pixel calibration of the image. This may be modified or created using the Calibration tool. |
| Image Orientation (Patient) | (0020,0037) | 1 | Image Plane | This field is required to use cross-sectional imaging tools. |
| Image Position (Patient) | (0020,0032) | 1 | Image Plane | This field is required to user cross-sectional imaging tools. |
| Conversion Type | (0008,0064) | 1 | SC Image | For Example “WSD” |
| Physical Units X Direction | (0018,6024) | 1 | US Frame of Reference | US Frame of Reference is used only when the unit is cm (0003H) |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Physical Units Y Direction | (0018,6026) | 1 | US Frame of Reference | US Frame of Reference is used only when the unit is cm (0003H) |
| Samples per Pixel | (0028,0002) | 1 | Image Pixel | Images can be displayed when this element is 1 or 3. |
| Photometric Interpretation | (0028,0004) | 1 | Image Pixel | Images can be displayed when this element is "MONOCHROME2", “MONOCHROME1”, "RGB", "YBR_FULL", "YBR_FULL_422", "YBR_PARTIAL_422", "YBR_ICT", "YBR_RCT" and "PALETTE COLOR". This element may be set to "MONOCHROME2", “MONOCHROME1”, "RGB", "YBR_FULL", "YBR_FULL_422", "YBR_RCT" and "YBR_ICT". |
| Rows | (0028,0010) | 1 | Image Pixel | No restriction is placed on the number of rows. |
| Columns | (0028,0011) | 1 | Image Pixel | No restriction is placed on the number of columns. |
| Ultrasound Color Data Present | (0028,0014) | 3 | US Image | This field is set for frame grabbed ultrasound color images. |
| Bits Allocated | (0028,0100) | 1 | Image Pixel | Images can be displayed when this field is set to between 1 bit to 16 bits. |
| Bits Stored | (0028,0101) | 1 | Image Pixel | Images can be displayed when this field is set to between 1 bit to 16 bits. |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| High Bit | (0028,0102) | 1 | Image Pixel | Images can be displayed when this field is set to between 1 to 15. |
| Pixel Representation | (0028,0103) | 1 | Image Pixel | This must be set to 0 or 1. |
| Planar Configuration | (0028,0006) | 1C | Image Pixel | Images can be displayed both color-by-plane and color-by-pixel. Created images always set this element to 1. |
| Red Palette Color Lookup Table Data | (0028,1201) | 1C | Image Pixel | Mandatory for images with a palette. This element is unused for created images as images with a palette are never created. |
| Green Palette Color Lookup Table Data | (0028,1202) | 1C | Image Pixel | Mandatory for images with a palette. |
| Blue Palette Color Lookup Table Data | (0028,1203) | 1C | Image Pixel | Mandatory for images with a palette. |
| Frame Time | (0018,1063) | 1C | Cine | This element is used to calculate the playback speed for multi-frame images. |
| Number of Frames | (0028,0008) | 1 | Multi-frame | This element is used to display multi-frame images. |
| Modality LUT Sequence | (0028,3000) | 1C | Modality LUT | Should be present if Rescale Intercept is not present for accurate display. |
| >LUT Descriptor | (0028,3002) | 1C | Modality LUT | Must contain three values describing the format of LUT Data. |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| >Modality LUT Type | (0028,3004) | 1C | Modality LUT | Specifies the units used in the LUT. |
| >LUT Data | (0028,3006) | 1C | Modality LUT | Contains the LUT entry values. |
| Rescale Intercept | (0028,1052) | 1C | Modality LUT | Should be present if Modality LUT Sequence is not present for accurate display. |
| Rescale Slope | (0028,1053) | 1C | Modality LUT | Should be present if Modality LUT Sequence is not present for accurate display. |
| Window Center | (0028,1050) | 3 | VOI LUT | Should be present for accurate display. |
| Window Width | (0028,1051) | 1C | VOI LUT | Should be present for accurate display. |

## Attribute Mapping

The mapping between received DICOM attributes and PowerReader Screns of PowerServer is shown in Table 9.1.3-1.

#### Table 9.1.3 -1 Attribute Mapping

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Patient's Name | (0010,0010) | Patient Info | Demographics | Last Name First Name Middle Name Prefix Suffix |
| Patient ID | (0010,0020) | Patient Info | Demographics | Patient ID |
| Issuer Of PatientID | (0010,0021) | Patient Info | Demographics | Issuer of Patient ID |
| Patient’s Birth Date | (0010,0030) | Patient Info | Demographics | Birth Date |
| Patient’s Sex | (0010,0040) | Patient Info | Demographics | Sex |
| Patient's Insurance | (0010,0050) | Patient Info | Demographics | Account Number |
| Plan Code | -> |  |  |  |
| Sequence -> Code | (0008,0100) |  |  |  |
| Value |  |  |  |  |

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Patient’s Primary Language Code Sequence -> Code Value | (0010,0101) -> (0008,0100) | Patient Info | Demographics | Language |
| Other Patient IDs | (0010,1000) | Patient Info | Demographics | SSN |
| Patient’s Mother’s Birth Name | (0010,1060) | Patient Info | Demographics | Mother’s Maiden Name |
| Ethnic Group | (0010,2160) | Patient Info | Demographics | Ethnicity |
| Race | (3129,1010) | Patient Info | Demographics | Race |
| Smoking Status | (0010,21A0) | Patient Info | Demographics | Smoking Status |
| Patient’s Address | (0010,1040) | Patient Info | Demographics | Address, Zip / Postal Code, City, State / Province |
| Country of Residence | (0010,2150) | Patient Info | Demographics | Country |
| Patient’s Telephone Numbers | (0010,2154) | Patient Info | Demographics | Home Phone Number, Business Phone Number, |

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Allergies | (0010,2110) | Patient Info | Health Information | Allergy Name Allergy Notes Severity |
| Additional Patient History | (0010,21B0) | Patient Info | Demographics | Additional Notes |
| Study Description | (0008,1030) | Study Info | Study Details | Study Description |
| StudyID | (0020,0010) | Study Info | Study Details | StudyID |
| Requested Procedure ID | (0040,1001) | Study Info | Study Details | Requested Procedure ID |
| Requested Procedure ID | (0040,1001) | Study Info | Study Details | Requested Procedure ID |
| Study Status ID | (0032,000A) | Study Info | Study Details | Status |
| Study Priority ID | (0032,000C) | Study Info | Study Details | Priority |
| Accession Number | (0008,0050) | Study Info | Study Details | Accession Number |
| Study Date | (0008,0020) | Study Info | Study Details | Date & Time |
| Study Time | (0008,0030) | Study Info | Study Details | Date & Time |

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Modalty | (0008,0060) | Study Info | Study Details | Modalty |
| Laterality | (0020,0062) | Study Info | Study Details | Laterality |
| Body Part Examined | (0018,0015) | Study Info | Study Details | Body Part |
| Institution Name | (0008,0080) | Study Info | Study Details | Imaging Facility |
| Institutional Department Name | (0008,1040) | Study Info | Study Details | Department |
| Current Patient Location | (0038,0300) | Study Info | Study Details | Patient’s Location |
| Scheduled Study Location | (0032,1020) | Study Info | Study Details | Exam Room |
| Procedure Codes Sequence -> Code Value | (0032,1064) -> (0008,0100) | Study Info | Procedure Codes | Procedure Codes |
| Procedure Codes Sequence -> Code Meaning | (0032,1064) -> (0008,0104) | Study Info | Procedure Codes | Description |
| Admitting Diagnoses Code Sequence -> Code Value | (0008,1084) -> (0008,0100) | Study Info | Diagnosis Codes | Diagnosis Code |

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Admitting Diagnoses Code Sequence -> Code Meaning | (0008,1084) -> (0008,0104) | Study Info | Procedure Codes | Description |
| Reason For Study | (0032,1030) | Study Info | Notes | Reason for Referral |
| Additional Patient History | (0010, 21B0) | Study Info | Notes | Clinical Notes |
| Study Comments | (0032,4000) | Study Info | Notes | Comments |
| Admission ID | (0038, 0010) | Study Info | Visit | Visit Number |
| Route of Admission | (0038, 0016) | Study Info | Visit | Patient Class |
| Visit Comments | (0038, 4000) | Study Info | Visit | Description |
| Ambulatory Status | (3135,1010) | Study Info | Visit | Ambulatory Status |
| Patient’s Height | (0010, 1020) | Study Info | Visit | Height (cm). |
| Patient’s Weight | (0010, 1030) | Study Info | Visit | Weight (kg). |
| Referring Physician’s Name | (0008,0090) | Study Info | Assigned Personal | Referring Physician |

| Entity Name | Tag | Screen of PowerReader | Tab or Sub Menu | Fields of PowerReader Screen |
| --- | --- | --- | --- | --- |
| Name Of Physician Reading Study | (0008,1060) | Study Info | Assigned Personal | Reading Physician |
| Performing Physician’s Name | (0008,1050) | Study Info | Assigned Personal | Performing Physician |
| Operator’s Name | (0008,1070) | Study Info | Assigned Personal | Performing Technologist |
| Interpretation Transcriber | (4008,010A) | Study Info | Assigned Personal | Transcriptionist |
| Physicians Of Record | (0008,1048) Value Multiplicity (1-n) | Study Info | Assigned Personal | List of Consulting Physicians. |

## Coerced / Modified Fields

Coercion and Modification of the fields can be performed using Custom Scripting Possibility for each received DICOM node. Custom Scripting Possibilitity is invoked from Station List (See Fig. 4.4.2-5)

s

## Data Dictionary of Private Attributes

#### Table 9.2-1

Data Dictionary of Private Attributes

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Filter Object Identification Code | (3109,0010) | 3 | Private Tag (LO) | Ramsoft Filter |
| Filter Object Sequence Group | (3109,1010) | 3 | Private Tag (SQ) | Filter Object Sequence Group |
| Filter Object Sequence Identification Code | (310B,0010) | 3 | Private Tag (LO) | Ramsoft Sequence Filter |
| Filter Object Sequence | (310B,1010) | 3 | Private Tag (SQ) | Filter Object Sequence |
| Filter Type | (310B,1020) | 3 | Private Tag (SS) | Possible values: Kernel Maximum Minimum Median Custom Values |
| Filter Size | (310B,1030) | 3 | Private Tag (SS) | Possible values: 3 X 3 5 X 5 7 X 7 |
| Filter Name | (310B,1040) | 3 | Private Tag (LO) | Possible values: Detect Edges Enhance Edges |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
|  |  |  |  | Maximum Minimum Median Shadowing Sharpening Smoothing Custom Value |
| Filter Kernel Sum | (310B,1050) | 3 | Private Tag (SL) | Sum of Filter Kernels Elements |
| Filter Kernel Elements | (310B,1060) | 3 | Private Tag (SL) (1-n) | Filter Kernel Element Value |
| Filter Modality | (310B,1070) | 3 | Private Tag (CS) | Filter Modality Code |
| Filter BodyPart | (310B,1080) | 3 | Private Tag (CS) | Filter BodyPart Code |
| Image Process Identification | (310D,0010) | 3 | Private Tag (LO) | Ramsoft Image Processing Identifier |
| Sharpness | (310D,1010) | 3 | Private Tag (DS) | Sharpness degree for DICOM PR object, range from 0.0 to 1.0 |
| Unsharp Mask | (310D,1020) | 3 | Private Tag (DS) | Unsharp mask degree for DICOM PR object, range from 0.0 to 1.0 |
| PS VOI LUT Function | (310D,1070) | 3 | Private Tag (CS) | VOI LUT function for DICOM PR object |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| PS VOI LUT Center | (310D,1080) | 3 | Private Tag (DS) | VOI LUT Window Center for DICOM PR object |
| PS VOI LUT Width | (310D,1090) | 3 | Private Tag (DS) | VOI LUT Window Width for DICOM PR object |
| Arbitrary Rotation Identification | (310F,0010) | 3 | Private Tag (LO) | RamSoft Arbitrary Rotation Identifier |
| Arbitrary Rotation | (310F,1010) | 3 | Private Tag (US) | Rotation angle for DICOM PR object. If rotation angle is oblique, it will be stored in this field. |
| RamSoft File Kind Identifier | (3111,0010) | 3 | Private Tag (LO) | RamSoft File Kind Identifier |
| File Kind | (3111,1010) | 3 | Private Tag (CS) | File kind (file extension) of DICOM SR object, e.g. DOCX, DOC, JPG, BMP. |
| Template File Kind | (3111,1020) | 3 | Private Tag (CS) | Template file kind (file extension) of DICOM SR object, e.g. DOT, DOTX, DOTM. |
| Custom Report Data Identification | (3113,0010) | 3 | Private Tag (LO) | RamSoft Custom Report Identifier |
| Custom Report Data | (3113,1010) | 3 | Private Tag (OB) | Custom report data in DICOM SR object |
| Custom Report Data Size | (3113,1020) | 3 | Private Tag (UL) | Size in bytes of custom report data in DICOM SR object |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Custom Report Template Data | (3113,1030) | 3 | Private Tag (OB) | Custom report Word template data in DICOM SR object |
| Custom Report Template Data Size | (3113,1040) | 3 | Private Tag (UL) | Size in bytes of custom report Word template data in DICOM SR object |
| Custom Object Reference Identification | (3115,0010) | 3 | Private Tag (LO) | RamSoft Custom Object Reference Identifier |
| Scout Reference UID | (3115,1010) | 3 | Private Tag (UI) | Scout Reference UID of DICOM PR object |
| Scout Reference Frame Number | (3115,1020) | 3 | Private Tag (US) | Scout Reference Frame Number of DICOM PR object |
| Ramsoft Color Mapping Identifier | (3117,0010) | 3 | Private Tag (LO) | Ramsoft Color Mapping Identifier |
| Color Mapping Name | (3117,1010) | 3 | Private Tag (CS) | Name of color mapping in DICOM PR object |
| Color Mapping Palette Size | (3117,1020) | 3 | Private Tag (US) | Palette size of color mapping in DICOM PR object |
| Color Mapping Palette | (3117,1030) | 3 | Private Tag (US) | Palette of color mapping in DICOM PR object |
| Document Type Identification | (311B,0010) | 3 | Private Tag (LO) | Document Type Identifier |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Document Type | (311B,1010) | 3 | Private Tag (LO) | Document type of DICOM SR object. Enumerated values: DIAGNOSTIC REPORT DIAGNOSTIC PRELIMINARY ADMIN CLINICAL STUDY FORM INSTRUCTION MAMMO RIS Rx SCREENING REFERRAL INSURANCE CARD PATIENT FORMS LABS PATIENT REGISTRATION FORM PATIENT INSURANCE CARDS HIPAA CONSENT FORM MEDICAL RECORDS RELEASE FORM PRIOR REPORT SURGERY CONSULTATION REPORT |
| RIS Page Number Identification | (311D,0010) | 3 | Private Tag (LO) | RIS Page Number Identifier |
| RIS Page Number | (311D,1010) | 3 | Private Tag (LO) | RIS Page Number Value |
| RamSoft DragonNS Data Identification | (311F,0010) | 3 | Private Tag (LO) | RamSoft DragonNS Data Identifier |
| Dragon NS Data | (311F,1010) | 3 | Private Tag (OB) | Dragon® Medical Data in DICOM SR object |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Dragon NS Data Size | (311F,1020) | 3 | Private Tag (UL) | Size in bytes of Dragon® Medical Data in DICOM SR Object |
| RamSoft Image Tissue Edge ROI Identification | (3121,0010) | 3 | Private Tag (LO) | RamSoft Image Tissue Edge ROI Identifier |
| Edge ROI | (3121,1010) | 3 | Private Tag (DS) | Edge ROI Value |
| Apply Auto Size of Edge ROI | (3121,1020) | 3 | Private Tag (CS) | Enumerated values: YES = Apply NO = Not Apply |
| Show Annotation Identificaion | (3123,0010) | 3 | Private Tag (LO) | Show Annotation Identifier |
| Show Annotation ROI | (3123,1010) | 3 | Private Tag (CS) | Show Annotation ROI Value |
| SUV Type Identification | (3125,0010) | 3 | Private Tag (LO) | SUV Type Identifier |
| SUV Type | (3125,1010) | 3 | Private Tag (SH) | SUV (Standardized Uptake Value) types of DICOM PR object. Enumerated values: SUVbw = body weight SUVlbm = lean body mass SUVbsa = body surface area |
| Document Protection Identification | (3127,0010) | 3 | Private Tag (LO) | Document Protection Identifier |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Document Protection Type | (3127,1010) | 3 | Private Tag (UL) | Document Protection Type of DICOM SR Object. Enumerated Values: 0xFFFFFFFF = No Protection 0x00000000 = Allow Only Revisions 0x00000001 = Allow Only Comments 0x00000002 = Allow Only Form Fields |
| Race | (3129,1010) | 3 | Private Tag (LO) | Enumerated values: 1002-5 (AMERICAN INDIAN OR ALASKA NATIVE) 2028-9 (ASIAN) 2054-5 (BLACK OR AFRICAN AMERICAN) 2076-8 (NATIVE HAWAIIAN OR OTHER PACIFIC ISLANDER) 2106-3 (WHITE) UNK (DECLINED) |
| Ambulatory Status Identification | (3135, 0010) | 3 |  |  |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Ambulatory Status | (3135,1010) | 3 | Private Tag (LO) | Enumerated values: A0 = NO FUNCTIONAL LIMITATIONS A1 = AMBULATES WITH ASSISTIVE DEVICE A2 = WHEELCHAIR/STRETCHER BOUND A3 = COMATOSE; NON-RESPONSIVE A4 = DISORIENTED A5 = VISION IMPAIRED A6 = HEARING IMPAIRED A7 = SPEECH IMPAIRED A8 = NON-ENGLISH SPEAKING A9 = FUNCTIONAL LEVEL UNKNOWN B1 = OXYGEN THERAPY B2 = SPECIAL EQUIPMENT (TUBES, IVS, CATHETERS) B3 = AMPUTEE B4 = MASTECTOMY B5 = PARAPLEGIC B6 = PREGNANT |
| Graphic Annotation InstanceUID | (3177,1010) | 3 | Private Tag (UI) | Graphic Annotation InstanceUID Code |
| Text Annotation Identification | (3179,0010) | 3 | Private Tag (LO) | Text Annotation Identifier |
| Referenced Graphic Annotation InstanceUID | (3179,1010) | 3 | Private Tag (UI) | Referenced Graphic Annotation InstanceUID Value |
| Show Text Annotation | (3179,1020) | 3 | Private Tag (CS) | Text Annotation Value |
| Fit To Width Identification | (3181,0010) | 3 | Private Tag (LO) | Fit To Width Identifier |
| Fit To Width | (3181,1010) | 3 | Private Tag (CS) | Fit To Width Value |

| Entity Name | Tag | Type | Module Name | Details |
| --- | --- | --- | --- | --- |
| Number List Preserve Identification | (3183,0010) | 3 | Private Tag (LO) | Number List Preserve Identifier |
| Number List Preserve | (3183,1010) | 3 | Private Tag (CS) | Number List Preserve Value |
| RamSoft Image Horizontal Flip Identification | (3185,0010) | 3 | Private Tag (LO) | RamSoft Image Horizontal Flip Identifier |
| RamSoft Image Horizontal Flip | (3185,1010) | 3 | Private Tag (CS) | RamSoft Image Horizontal Flip Value |
| RamSoft Image Rotation Identification | (3187,0010) | 3 | Private Tag (LO) | RamSoft Image Horizontal Flip Identifier |
| RamSoft Image Rotation | (3187,1010) | 3 | Private Tag (CS) | RamSoft Image Rotation Value |

## DICOM Patient, Visit, Study and Series Fields stored in DB

The following DICOM Patient, Visit, Study and Series Fields are stored in the database:

- Patient ID (0010, 0020)

- Issuer of Patient ID (0010, 0021)

- Patient Name (0010, 0010)

- Patient Birth Date (0010, 0030)

- Patient Sex (0010, 0040)

- Patient Address (0010, 1040)

- Patient Country of Residence (0010, 2150)

- Patient Telephone Numbers (0010, 2154)

- Patient Telecom Info (0010, 2155)

- Patient Primary Language (0010, 0101)

- Patient Insurance Plan (0010, 0050)

- Other Patient IDs (0010, 1000)

- Patient Comments (0010, 4000)

- Patient Contrast Allergies (0010, 2110)

- Patient's Weight (0010, 1030)

- Patient's Size (0010, 1020)

- Patient's Race (3129, 1010)

- Patient's Ethnic Group (0010, 2160)

- Patient's smoking status (0010, 21A0)

- Patient's mother birth name (0010, 1060)

- Admission ID (Visit) (0038, 0010)

- Admitting Date (Visit) (0038, 0020)

- Admitting Time (Visit) (0038, 0021)

- Visit Comments (0038, 4000)

- Route of Admissions (0038, 0016)

- Ambulatory Status (3135, 1010)

- Study Instance UID (0020, 000d)

- Study Priority ID (0032, 000C)

- Requested Procedure Priority (0040, 1003)

- Study Date (0008, 0020)

- Study Time (0008, 0030)

- Study Description (0008, 1030)

- Study ID (0020, 0010)

- Accession Number (0008, 0050)

- Requested Procedure ID (0040, 1001)

- Reading Physician Name (0008, 1060)

- Referring Physician Name $0008, 0090)

- Performing Physician Name (0008, 1050)

- Operator Name (Technologist) (0008, 1070)

- Transcriptionist Name (4008, 010A)

- Procedure Code Sequence (0008, 1032)

- Diagnosis Code Sequence (0008, 1084)

- Study Status ID (0032, 000A)

- Institution Name (0008, 0080)

- Study Comments (0032, 4000)

- Current Patient Location (0038, 0300)

- Reason for Study (0032, 1030)

- Reason for the Imaging Service Request (0040, 2001)

- Additional Patient History (0010, 21b0)

- Scheduled Study Location (0032, 1020)

- Patient's Institution Residence (0038, 0400)

- Institutional Department Name (0008, 1040)

- Physician's of Record (Consulting Physicians) (0008, 1048)

- Series Date (0008, 0021)

- Series Time (0008, 0031)

- Series Description (0008, 103e)

- Series Instance UID (0020, 000e)

- Series Number (0020, 0011)

- Series Laterality (0020, 0060)

- Body Part Examined (0018, 0015)

- Modality (0008, 0060)

## Revision History

| Version | Revision Notes | Updated by | Reviewed by | Effective Date |
| --- | --- | --- | --- | --- |
| Rev. A | Initial version for 6.0 | VKM | VRN | 2015/05/28 |
| Rev. B | TLS conformance updated | VRN | VRN | 2015/07/06 |
| Rev. C | Added Document DICOM Patient, Visit, Study and Series Fields stored in DB | JZG |  |  |
| Rev. D | - Add Negotiation Fuji Computed Radiography Image Storage. - TLS conformance updated - Add DICOM WEB Support for Client Echo, Store, Find, Retrieve operations | VKM | LLG | 2020/06/28 |
| Rev. E | - Add Negotiation for Ophthalmic Tomography Image | SVN | LLG | 2020/12/31 |
| Rev. F | Added DIMSE Query/Retrieve (C-FIND, C-MOVE, C-GET) Transfer Syntax Support subsections in Sections 4.2.7–4.2.11; added Storage SOP Class no-change note in Sections 4.2.3.1 and 4.2.4.1. | SD |  | 2026/04/16 |